## Slide 1: T-REX Toll Gate 2 Time-of-flight Reciprocal space EXplorer
Lund, 8th March 2017
A bispectral chopper spectrometer for
magnetism, material science and soft matter
Nicolò Violini, Jörg Voigt, Earl Babcock, Zahir Sahli, Günter Kemmerling, Klaus Bussmann, Thomas Brückel
Hans Kämmerling,
Achim Heynen
Andrea Orecchini, Alessandro Paciaroni, Marco Zanatta, Pietro Tozzi, Francesco Sacchetti

## Slide 2
ConOps / T-REX Scientific Scope
Overview of Science Case
“project” vs “full scope” after SSM
Overview of T-REX layout & PBS & P&ID
      upgrade stages from IOSP doc
      QA COnOps and IOSP-SP
SysReq / T-REX HLSR in connection to Scientific Scope
  Functional requirements for High Level Components
  QA SysReq
PSDD / Description of Preliminary Design
  Interfaces with monolith and bunker and neighbors
  Key components and areas
  Performance at 2 MW
QA PSDD
WPS + IOSP / Instrument life-cycle
  Schedule
  Risk analysis
  QA WPS and IOSP-IO
Outline of TG2 presentation
Address issues from reviews

## Slide 3
A spectrometer for a broad user community
Wide energy and Q range
Instrument configurations
energy resolution
Q resolution
polarization analysis
Sample environment
Predict trends and developments
3
30%-40%
20-30% +
= more chances
of early success
SNS
ISIS
5D mapping

## Slide 4: Order at high temperature High energy excitations
Interplay lattice, charge, spin  Distinguish excitations
Resolving gaps and zone boundary behavior
Line shapes
Identify spurions with PA
Hour glass magnetic spectrum in La2-xSrxCoO4.
ILL Science highlight
Y. Drees et al., Nat. Comm. 4:2449
Thermal energy range (beyond 100 meV) indispensable!
90
85
80
75
35
30
25
20
25
20
15
10
5

## Slide 5: Functional Materials: Multiferroics
IN5: Ba3NbFe3Si2O14
Loire et al., PRL 106, 207201 (2011)
Very good energy resolution
Good Q resolution
Test ground for theory (line shape …)

## Slide 6
Very good energy resolution
Good Q resolution
Test ground for theory (line shape …)
IN5: Ba3NbFe3Si2O14
Loire et al., PRL 106, 207201 (2011)
PA: now TAS; future: T-REX
High Intensity, High Resolution
&
Polarization Analysis
Functional Materials: Multiferroics

## Slide 7
W. Paulus et al, JACS 130, (2008), 16080
Thermoelectrics: Electron-Phonon coupling: Line shape at high energy
Ion Transport: QENS at large momentum transfer + PA (H, Na, Li, Cl, N, …)
From lattice dynamics
to diffusive motion
Thermal energies to tens of µeV
+
large Q
Olivier Delaire et al. PNAS 2011;108:4725-4730
Functional Materials

## Slide 8: Soft Matter & Life Science
Dynamics in bio-molecular components:
Proteins, DNA, RNA and lipid membranes are central to cellular function.
Biological functions are governed by their dynamical character, particularly on the picosecond time-scale.
Diffusional and relaxational motions
Widely studied by QENS
T-REX:
  Cold Ei and E down to 20 eV
  PA
Vibrational collective motions.
…Purely Inelastic Scattering
New topic
Thermal Ei
Flux limited
Frequency (cm-1)
Crystal Angle (degrees)

## Slide 9
INS close to forward direction
collective coherent
excitations in glasses
Zanatta et al.,
J. Phys. Chem. Lett. 2013, 4, 1143−1147
Tof-Tof
BRISP

## Slide 10: T-REX “Full scope” Science Cloud

## Slide 11: “Project scope” agreed at SSM

## Slide 12: Project scope vs Full scope as agreed at SSM
It was agreed that the instrument presented as configuration 2 in the scope-setting report will form the basis of the scope and budget for T-REX. It will be world-leading at 2MW beam power when it comes online by an order of magnitude or more and has a technically straightforward upgrade path to full scope.
Chart: E_i max (meV)
| 80.0 | 300.0 | 80.0 | 60.0 | 2000.0 | 1500.0 | 2000.0 | 2000.0 | 30.0 | 100.0 | 25.0 | 4.87
instrument | 10.1786 | 16.101 | 11.6765 | 9.9419 | 31.0646 | 49.7097 | 31.0646 | 57.3998 | 7.1503 | 12.0313 | 6.4175 | 2.5857
50% | 12.4258 | 1.3892
cost cat | 5.1378 | 0.57442
full scope | 16.7127 | 1.8685
5D mapping
Polychromatic experiments
+ PA unique feature

## Slide 13: T-REX layout project scope
D03
E02
E01

## Slide 14: T-REX Product Breakdown Structure
ESS standard nomenclature
11 High Level components
Nearly 200 Low Levels components
Baseline structure for SysReq and Budget
High resolution pdf available in Indico

## Slide 15: PBS – a few more details
Description of components follows in section ‘key components & areas’

## Slide 16: PBS – a few more details
Description of components follows in section ‘key components & areas’

## Slide 17
High resolution pdf available in Indico
network
control
hardware
media

## Slide 18
T-REX life cycle (wrong numbering of TGRs)

## Slide 19
T-REX life cycle (correct numbering of TGRs)

## Slide 20: Budget

## Slide 21
Upgrade stages to full scope
Increase Detectors coverage up to 2.2 sr
Sample Environment Equipment
+T0C
4-P choppers
Not mentioned in IOSP document
Faster choppers
Supermirror analyser
+2.3 M€ MG
+4.9 M€ 3He
+1.667 M€
+240 k€
~ 0.5M€
~ M€
~ M€
2022
2022
2024
When available
2025
PROPOSED START
2022
Chart: Budget (M€)
2022 | 1.8
2023 | 1.6
2024 | 1.4
2025 | 2

## Slide 22
Sample Environment Priorities (Cost + ESS pool based)
CCR
ILL Furnace
Clamp cells < 3 GPa
6kV HV supply
3He Sorption Stick
Humidity Chamber
Vertical cryomagnet 7T (1) Paris-Edinburgh Cell (1)
Gas cells < 1 GPa (2)
Gas Handling (2)
+1200k€
Project scope
Full Scope
IR furnace (1)
ES Levitator
Pump & Probe set-up
Future upgrade
Rely on ESS-pool on day 1
Top-down, side access possible
Magic PASTIS < 1m diameter
Sample area shall enable installation of XL SEE
<1000 kg, height < 1.7 m, d < 0.8 m
Orange cryofurnace
+219k€
+250k€

## Slide 23: More Questions about ConOps and IOSP (Staging Plan) ?

## Slide 24
Science driven High Level Requirements / Design Goals
1 | Diffusion and relaxations High Tc SC/materials | Incident Energy Range | 2 meV - 160 meV
2 | small gaps/ lineshape/QENS Magnetism/materials | Elastic energy resolution (FWHM) | 1% at 2 meV 3% at 100 meV
3 | Ion transport/phonons/ diffusion | Q range | 0.05 Å-1 – 10 Å-1 0.05 Å-1 – 15 Å-1 (full scope)
4 | Single x-tals | Q resolution (FWHM) | 0.05 Å-1 at 80 meV tailored to needs
5 | Magnetic s. Factor/ disorder | Minimum 2ϑS | 1°
6 | Single x-tals + powders | Beam size | 10mmx30mm
7 | Single x-tals | Beam profile | flat
8 | Magnetism/Single x-tals | Polarization Analysis | X-Y-Z
9 |  | S/N | maximum

## Slide 25: S/N requirement (not in SysReq doc)
From M. Arai IKON 12 presentation
Energy or time domain
Which subframe? Ei ≥ 50 meV
Which TOF? Beyond V and Al phonon cut-off
Sample conditions? 10% V Cold to reduce energy up-scattering
Consider scattering from SEE
Sample inherent background
Better discrimination of ‘unwanted’ features:
PA and large angle detectors
Taking Signal from Vitess/McStas simulations S= 5·104 n/s
Noise from MCNP simulations for MAGIC
N= 4 n/s/cm2 x 100 cm2 x 0.01 (detection efficiency)
S/N 104 expected !!no validation in project scope!!
?? Beam stop, skyshine, groundshine, cross talks, neutronics calculations ??
HRC

## Slide 26
Functional Requirements
Beam Transport and Conditioning System
n’s wavelength transport: 0.7Å – 6.5 Å
band selection (continuous): Δλ = 1.67Å
suppression (rate): others to 10-4
resolution: 0.5% < δλ/λ < 3%
ToF resolution selection 10 µs < δt/t < 80 µs
Beam size 10mm x 30mm (W x H)
tailor to sample
profile flat
collimation λ dependent from +/- 0.25° to +/-1° improvable +/- 0.25°, +/-0.5°
profile smooth
polarization flipping ratio shall ≥ 20
monitors time resolution of 2 µs
S/N breaking the LoS brilliance transfer should be 100%

## Slide 27
Functional Requirements
Sample Exposure System
Sample positioning: a ± 0.1 mm, p ± 0.1 mm
rotation: a ± 0.5°, p ± 0.1°
Sample Environment Equipment XL SEE
diameter < 1 m
weight < 2 t
4 m free height from sample
Interface with Instrument Control

## Slide 28
Sample-detector distance: 3 m ± 1 mm
Large solid angle: vertical -15° to +25°
horizontal +1° to +70° (project scope)
Minimum angle: 1°
Time resolution (FWHM): ≤ 5 µs
PSD Positional resolution (FWHM): 20 mm x 20 mm x 10 mm
Detection Efficiency: ≥ 70 % at 1.8 Å
average 74 %, 1.6 Å < λ < 5 Å
Count Rate: <104 Hz/cm2
Neutron Selectivity: <10-6 gamma detection prob.
Noise: <10-4 events/s/cm2
Vacuum ≤ 10-6 mbar
Prevention of secondary scattering
Scattered beam polarization flipping ratio shall ≥ 20
~ 20 m2
Functional Requirements
Scattering Characterization System
-36° to +144° (full scope)

## Slide 29
Wavelength rejection: others
ToF resolution achieve 10 µs (M choppers)  max frequency
100 µs (P choppers)  max frequency
selection: 10 µs - 80 µs  given LM/LP
Frame overlap avoidance
Frequency stability, precision
Position precision
Vacuum HV
Interface with the other BTS sub-systems ( beam quality)
with the other systems and subsystems ( materials)
+ Non functional requirements:
Reliability Accessibility Maintenance Integration
Safety
Functional Requirements
Beam Transport and Conditioning System
Choppers (not requested in SysReq doc)

## Slide 30: High Level Scientific Requirements
The instrument shall allow data to be collected to minimum energy transfer of 20 µeV.
The instrument shall allow data to be collected to maximum energy transfer of 140 meV.
The instrument shall allow data to be collected to minimum wavevector transfer of 0.05 Å-1.
The instrument shall allow data to be collected to maximum wavevector transfer of 10 Å-1.
The instrument shall match the size of the neutron beam to the size of the sample, assuring the homogenous illumination of rotating samples.
The instrument shall allow data to be collected to minimum scattering angle of 1°.
The instrument shall allow the analysis of neutron spin polarization.
The instrument shall allow the energy resolution to be optimized for the experiment.
The instrument should allow the Q resolution to be optimized for the experiment.
The instrument should be capable of providing elastic energy resolution of 1% at incident energy of 2 meV.
The instrument should be capable of providing energy resolution of 3% for an energy transfer of 100 meV.
The instrument should be capable of providing a Q resolution in elastic scattering condition of 0.05 Å-1 at incident energy 80 meV.
The instrument should allow data collection from samples < 102 mm3 volume.
The instrument should maximize the signal-to-background (S/B) ratio in the scattering pattern.
The instrument shall allow the control of the physical and thermodynamical condition of the sample.
The System’s design shall provide the space and flexibility necessary to host and drive the staging process to achieve the full scientific scope.
T-REX should serve the user, science and instrumental development program without interruptions during source operation.

## Slide 31: More Questions about System Requirements Doc?

## Slide 32
E01
Proton Beam
Bunker
Monolith
E02
D03
Layout

## Slide 33: Neighbours E01

## Slide 34: Shielding footprint D03
W7 T-REX
W6 MAGIC
W8 HEIMDAL
D03

## Slide 35
Layout – D03
Monolith, Bunker, D03 with Bandwidth Choppers BWC 1 and BWC 2
BWC 1
at 32 m
BWC 2 at 40 m

## Slide 36
2D view monolith + D03

## Slide 37
Target- monolith interface
Monolith: Cut at upper Moderator Center
Proton Beam
W7 T–REX
Upper Moderator

## Slide 38
Target- monolith interface
Target Wheel
Upper
Moderator
T –REX
centreline

## Slide 39
In-monolith insert (not in project scope)
NBOA Drawings delivered to ESS in Jan ‘17
Neutron Beam Optics Assembly | 
Entrance | 30 mm x 90 m (HxW)
Exit | 46 mm x 60 mm (HxW)
Thickness of the walls | 8 mm
Material | Al or Cu

## Slide 40: Light shutter system
Integrates the Cold Neutrons Extraction system
Drawings and model ‘ESS-0092689’
200 mm

## Slide 41: Light shutter system
Integrates the Cold Neutrons Extraction system
Drawings and model ‘ESS-0092689’
braked
vertical
servo
50 µm
pos resolution
200 mm
Guide
Bender + guide
Polarizing bender
+ guide
Saturation field
1 kG

## Slide 42: Neighbours inside bunker
Proposed position of T0 chopper
Proposed position of heavy shutter
Close collaboration and follow up with neighbours and CG

## Slide 43: T0 chopper (not in project scope)
Phase 1 study for DREAM (Neutronics + engineering)
ESS Chopper group:
Prototype development envisaged possibly at FZJ, managed by Dariusz Zielinski
Proposal of provision in project scope:
Make easier later installation/alignment inside bunker
Removable guide section of equal-size
Bottom part of the CHIM, +50k€
Further development for installation outside bunker (may not fit into neutron guide 60 mm x 85 mm WxH)
To be installed if operational experience shows that it is necessary

## Slide 44: Heavy shutter inside bunker
From Gàbor Laszlo – IKON12 Feb 2017

## Slide 45: Bunker feed-thru
Neutron Optics Drawings delivered to ESS
in Feb 2017  bunker procurement
Guide Support and alignment tbd

## Slide 46: BW choppers + shielding D03
Access to choppers
Horizontal split CHIM
load limit 14 t/m2
32 m
40 m
28 m
45 m

## Slide 47
E01
Proton Beam
Bunker
Monolith
E02
D03
Layout
Notes: comment e02
next slides e01

## Slide 48: 3D view of T-REX in E01

## Slide 49: T-REX (view from E01 – without exp. cave)

## Slide 50: Around the sample area
FAN chopper
Adiabatic Field rotation
Primary beam Collimators
Flange
Vessel/chopper
Guide field
x-y-z
Neutron beam
centreline
sample
MAGIC
PASTIS
M chopper

## Slide 51: PBS – a few more details (a remind of slides 15-16)
Neutron guide
Choppers
Collimators/slits
Polarization Equipment
Monitors
Shielding

## Slide 52
insert axis -17 mm
From back tracing option
Neutron guide (baseline) from “Report on Neutron Optics”
Notes: LoS
axis
piling corridor
ESS standard
robust design
m-index optimal for cost and profile

## Slide 53
Substrates (baseline) from “Report on Neutron Optics”
borofloat
Borkron N-BK7
metallic
100 years lifetime
Neutron losses
Needs revision after
new document ‘ESS-0097645’
metallic
Borkron N-BK7
OK with ‘ESS-0039408’
Notes: vitess simulation give order of magnitude of neutron losses
substrates lifetime
met 10^19
bk7 10^17
borofloat 10^16

## Slide 54: Brilliance transfer (Vitess calculations) from “Report on Neutron Optics”
Current design top moderator
It looks worse than it is!
30 mm beam height challenging with 30 mm moderator
Optimal for collimated thermal beam – Q resolution
Using focus-to-focus
Same flux expected from 6cm moderator (intrinsic gain) for cold neutrons
McStas calculations to validate

## Slide 55
Neutron guide (alternatives 1)
from “Report on Neutron Optics”
Breaking LoS earlier may reduce shielding cost
Breaking LoS at bunker wall costs 1 order of magnitude at 100 meV n. energy

## Slide 56
Neutron guide (alternatives 2)
from “Report on Neutron Optics”
Breaking LoS with vertical kink
Supermirrors m=6 needed
Low performance expected
MCNP calculations needed moderator emission and background at sample

## Slide 57
Insert axis
2000
5500
6000
curve
45
60
optics axis -17mm
shutter
light
12
Cold Neutrons Extraction System
Radius of curvature 7.16 m
400 Si wafers of 150 µm
m=4 supermirror
L=50mm
Front view
60 mm
46 mm
0.4°
Notes: butterfly 1 vs butterfly 2
minor changes
accept lower performance in cold range 1-2 y

## Slide 58
Expected cold-n extraction performance
2D POS
2D DIV
2Å
Vitess

## Slide 59
Choppers overview
Notes: Lp Lm
PC 4 slits
MC 2 slits

## Slide 60: Choppers Group Review
T0 chopper outside bunker
FAN chopper
FAN chopper development
M chopper 336Hz challenging
Magnetic field limits at the M chopper
PBS and WBS and budget for provision of T0 chopper
Early procurement is strongly advised
Notes: restrict presentation to critical aspects
development for T0 outside
development of FAN integration into ESS
lifetime of carbon fiber disks

## Slide 61: FAN chopper enables QENS @ ps
M-chopper
Fan chopper
170
165
160
Umbra shadows > tM (~20 µs)
Penumbra shadows < 2/fM (< 5.95 ms)
 5º < blade size < 30º , large freedom for device layout
14 Hz

## Slide 62: First tests at LET @ ISIS
Proof of principle with a prototype not designed for LET
Issues:
Synchronization with source
Encoders
Blade width not adapted
Change synchronization scheme
Test in Jülich with Stroboscope
From Rexroth to Beckhoff control
Refine mechanical setup
Instrument specific device compatible with ESS standards
Follow-up/communicate with ESS CG
1 2 3 45 678

## Slide 63: Primary beam collimators
Clean up the divergence profiles
Vacuum compatible
Guide field compatible
2D collimation with honeycomb
or 1D collimation with soller
Al foils 0.1mm Gd paint
99 cm
NIMA 489, 304 (2002)

## Slide 64: Beam shaping slits
Reduce beam size < 10 mm x 30 mm
Compact (along the beam)
Vacuum compatible
Magnetic field compatible
Al frame
Ceramics tracks
Piezo actuators
B4C blades

## Slide 65
Polarization Equipment of T-REX
3He cell
Magic PASTIS
guide field along the guide (20 Gauss)
Polarizing bender

## Slide 66: Guide field
Inside bunker
Power supply
Construction cost and schedule to be investigated
Or other permanent magnets?
NdFeB -20-30% magnetization in 2 years
(see MAGIC simulations MCNP)
Outside bunker
Soft
≥ 200 mm
60 mm
85 mm
Al Vacuum casing 10mm
Al/Cu substrate 8mm
NdFeB
magnets
Iron yoke
Al coils
j=2A/mm2
Soft iron sheet
5 mm
potential alternative
risk mitigation drawings of bunker feedthru
+40 mm from centerline 106 n/s/cm2

## Slide 67: Exchange Unit Neutron Guide / Polarizer
Neutron Guide
Polarizer
Development for TOPAS
Construction drawings available
Minor modifications needed

## Slide 68
Laser System
Mirrors
Magicbox
Oven
3He-Cell
Thermal neutrons Polarizer
TOPAS
Includes Adiabatic Fast Passage
Does not need 3He supply

## Slide 69
Adiabatic Field Rotation
between Fan Chopper and M Chopper Larger distance and less feedback to/on PASTIS
X or Y or Z guide field from M chopper to sample
Rotation of Hallbach magnet rings for transverse directions + solenoid for longitudinal
Direct Double Helix (DDH™) around beam tubes/collimators
Rotation around [111] direction of Hallbach magnet array
Adiabatic rotation

## Slide 70: MAGIC PASTIS
X-Y-Z analysis
Development for TOPAS
Construction drawings available
Needs 3He support lab/staff to provide 3He cells, needs set of cells
Reduce blind spots/
adapt to detectors gaps /
Adapt to detector height

## Slide 71: Load Lock system / sample change
Independent vacuum system/environment for sample change
15 min recovery 10-4 mbar
Common vacuum during measurements
Last window at M chopper
Simple

## Slide 72: Neutron Beam Monitors
Monitor 1
Monitor 2
Monitor 3
CDT Design
Vacuum compatible
<15mm thickness
2D readout
Gas supply (depends on provider 3He, Ar, CO2,N2)

## Slide 73: Shielding outside the bunker
Outside LoS
Gammas from th. n’s capture Ti
< 14 t /m2
Present layout exceeds
5 t /m2
< 20 t /m2
MCNP model of T-REX available
Calculations are CPU demanding, source-code license not available yet
Use of the ‘thermsource’ (spectra at 2m) is envisaged
Access to SC possible
Within LoS
Fast neutrons / Spallation
Vitess: 103-104 n/s/cm2 on substrate 0.1-10Å - 20cm concrete
MCNP MAGIC: 104n/s 1eV-1GeV – 50 cm concrete

## Slide 74: Shielding Experimental Cave
1 μSv/h
1 μSv/h
0.2 μSv/h
0.3 m
concrete
160 μSv/h
8 μSv/h
18 μSv/h
10.5 m
8 m
11 m
2.3 m
No direct view 
thermal flux
≤109 n/s converted in gammas
by 3 cm3 FeCd
6 m high walls
Load = 4.5 t/m2

## Slide 75: From IKON12 – M. Arai
Previous experience has a value
Main hall @ J-PARC

## Slide 76: From IKON12 – M. Arai

## Slide 77
From IKON12 – M. Arai
High energy neutrons eliminated either by T0 chopper or a curved guide
Ex. Los 28m, 40m instrument
T0 and curved guide are not perfect  use their combination in full scope
Gain from larger distance to source
Longer source pulse ~ 10 ms prompt flash affects 14meV<Ei<28meV

## Slide 78
PBS – a few more details (a remind of slides 15-16)
Detector Vessel
Detectors
Get Lost tube Beam Stop
Radial collimator

## Slide 79
Detector Vessel
10-6 mbar
Stainless steel 10 mm
Magnetic req.
2m R from sample
Access from lateral door
Inside:
Primary collimator,
Holding field,
PASTIS / SEE,
Radial collimators,
Detectors,
Get lost tube, beam stop

## Slide 80: Vacuum system for Detector Vessel
pumps
4 roots
3 turbo
1 screw
1 cryo
< 2 h to restore

## Slide 81
Multi-GRID detectors (baseline)
Prototype under construction
20-24 10B4C layers
6 x 1 µm, 10 x 1.25 µm, 4 x 2 µm
to reach 45-48% efficiency @ 1Å
Fairly compares to 3He at 6 bar
Pure Al frames
MG @ CNCS
Size = half of “8-pack” module
Installation June-July 2016
Test at spectrometer
Operation for 6+ months
Side-by-side comparison to He3
User experiments

## Slide 82
1mm gap
2mm gap
20 columns
96 rows
Multi-GRID detectors (baseline)
Full scope needs
100 modules = 5 banks x 20 modules
Project scope
40 modules = 2 banks x 20 modules
CNCS
prototype

## Slide 83
1 bank = 1 box
AlMg4.5Mn
Thickness of the Walls
  Sides 7 mm
  Top 10 mm
  Bottom 10 mm
  Back 5 mm
  Front 2 mm
  Ribs 2 mm
36°
Inner Radius 3000 mm
2200 mm
sample
Multi-GRID detectors (baseline)

## Slide 84
Deformation
< 1.5 mm
Stress
< 75N/mm2
Pressure
500 mbar
Multi-GRID detectors (baseline)

## Slide 85
3He detectors (alternative)
TOPAS
Design
Alternative within budget and minimum scientific impact
Full scope:
368 1” PSD tubes, 2.2 m high
23 x ‘16 packs’
Project scope (based on todays price information)
144 tubes, 9 packs
6 bars
Construction drawings available for mechanics
Vacuum compatible
Potential issue with detection rate req. >400KHz/tube from 10% sc. Vanadium
Needs new construction or tubes assembly

## Slide 86: Get-lost tube / Beam stop
3m
MAGIC
presentation at TG2
Neutronics calculations needed to define materials and dimensions
2x
Cave shielding
4 Å data
Normal guide
with primary
collimator

## Slide 87
Radial Oscillating collimator
Rin = 50 cm
Rout= 70 cm
Vertical -25°to +15 °
Horizontal -149°to +41°
Full scope detector coverage
Kapton foils 50 µm
+50 µm
Gd2O3 or 10B paint
1 °separation
Oscillation
+/- 5 °
0.1 Hz

## Slide 88: Performance @2MW
2 anticipated choppers configurations:
High Flux
5·106 n/s/cm2 @ 10 meV
2·106 n/s/cm2 @ 60 meV
Calculation method:
‘Vitess OT6’ x gain factors x B.T. Ratio x 2/5
High Resolution
3% el en resolution at Ei 100 meV
1% el en resolution at Ei 2 meV
++PA available on day 1
50% detector coverage is
less critical for thermal neutrons
Butterfly 2 moderator is more critical for cold neutrons
LET
CNCS
IN5
AMATERAS
4-SEASONS
3x IN5 same config.
6x 4-SEASONS same config.
x RRM

## Slide 89
Cold-n energy
HR ( E , Q )
4D INS
Spin correlations
Q-range is the limiting factor
Baker, Nat Phys (2012)
Cr8 rings
Detection area vs scientific impact
making qbits
Chart: E_i max (meV)
| 80.0 | 300.0 | 80.0 | 60.0 | 2000.0 | 1500.0 | 2000.0 | 2000.0 | 30.0 | 100.0 | 25.0 | 4.87
instrument | 10.1786 | 16.101 | 11.6765 | 9.9419 | 31.0646 | 49.7097 | 31.0646 | 57.3998 | 7.1503 | 12.0313 | 6.4175 | 2.5857
50% | 12.4258 | 1.3892
cost cat | 5.1378 | 0.57442
full scope | 16.7127 | 1.8685
IN5 equivalent resolution at longer wavelength

## Slide 90
TOPOLOGY
T-REX first day
Topology
Condensed Matter
Materials Science
Life Science
Superconductors
Insulators transitions
Higgs Mechanism
Quantum phase transitions
Formation of defects
Crystals
Ion-conductivity
nanostructured
materials
Entanglement
Human telomere
sequences
DNA, RNA
conformation
polymers
Oriented
Crystals
of proteins
thermoelectrics
Inter-chain interactions
Low neutron energy
K. Bera et al.,
PRB 91, 144414
Highlight: Haldane chains using Multiple energy scales
Intra-chain interactions:
High neutron energy

## Slide 91: More Questions about Preliminary System Description?

## Slide 92
T-REX life cycle

## Slide 93
T-REX schedule available (MS project format)

## Slide 94
Tailoring installation to access plan
Inserts to be installed Mar-’19
E01 - access
Apr-’19
In bunker
Nov ‘19-May ’20
From cave
back towards bunker

## Slide 96
Critical dates/MS of T-REX project
Detector Vessel -- Early procurement necessary --
Q2/2017 Preliminary Engineering design and interfaces
Q3/2017 Final design (WPS progress MS 6 Mar-18)
Q2/2019 Delivery at ESS (WPS progress MS 17 Apr-20)
Q4/2019 Commissioning
Detectors -- Early procurement? --
Q2/2017 MG prototype available for tests
Q1/2018 engineering design
Choppers -- Early procurement advised --
Q1/2018 engineering design
Q1/2018 contract signed
Neutron Optics inside bunker -- Early procurement under consideration --
Q4/2017 engineering design
Q2/2018 contract signed
Neutron Optics outside bunker & Shielding
Q2/2018 engineering design
Q4/2018 contract signed

## Slide 97
 | T -REX | TOPAS
Material | Stainless steel | Stainless steel
Vacuum | < 10-5 mbar | < 10-5 mbar
Height | 3250 mm | 3200 mm
Diameter | 8500 mm | 6500 mm
Volume | 85 m3 | 75 m3
Horizontal det. angles | -36°- 144° | -30° - 140°
Detector Vessel : Schedule & Experience

## Slide 98
Vessel delivered
6 months later at the earliest
Timeline Detector Vessel – scenario 1 manufacturing at FZJ
Access to E01

## Slide 99: Timeline Detector Vessel – scenario 2 external procurement
Need for early procurement independent from scenarios

## Slide 100
Finish 06/22
Half 2, 2019
Half 1, 2020
Half 2, 2020
Half 1, 2021
Half 2, 2021
Assembly
06/20 - 07/21
Coating production
10/19 - 05/20
Procurement
10/19 - 02/20
Installation
09/20 - 10/21
Commissioning
06/21 - 06/22
Detector Vessel is delivered at ESS
07/20
Installation Readiness Review 05/20
Timeline Detectors : MG production (ESS)
Scenario:
T-REX detectors manufactured after completion of C-SPEC detectors (project scope)
Half 1, 2022

## Slide 101
Timeline Choppers
Q2 2017
Q3 2019
Design
Construction
Assembling
Test
Clearly need for early procurement

## Slide 102
Start 01/17
Finish 02/21
Mar '17
Jun '17
Sep '17
Dec '17
Mar '18
Jun '18
Sep '18
Dec '18
Mar '19
Jun '19
Sep '19
Dec '19
Mar '20
Jun '20
Sep '20
Dec '20
Neutronics Final Design 01/17 - 09/17
Engineering Drawings 10/17 - 03/18
Engineer Drawings of mechanical structure 03/18 - 06/18
Procurement 06/18 - 12/18
Production 01/19 - 01/20
Assembly & alignment 03/20 - 06/20
Conceptual Design 01/17 - 06/17
MCNP final simulations 07/17 - 10/17
Report 10/17 - 11/17
Procurement 09/18 - 03/19
Construction 04/19 - 07/20
transport to Lund 07/20 - 08/20
installation 08/20 - 10/20
Engineering Drawings 11/17 - 09/18
transport to Lund 01/20 - 02/20
optics
shielding
Timeline Optics and Shielding
Access to D03

## Slide 103: Risk Analysis / technical risks

## Slide 104: Risk Analysis / cost risks

## Slide 105: Risk Analysis / schedule risks

## Slide 106: Hot commissioning draft: validation of requirements
Validation of PSS and dose levels
Neutron Beam validation (needs a plan with/from NOSG):
Validate beam characteristics upstream\downstream choppers
At sample: validate beam with monitors, validate divergence with rocking curve of crystal
Fine alignment may be needed afterwards
Calibration with standard samples V,
measure S/N
Measure first inelastic features from standard samples
Validate flipping ratio:
Cold neutrons: bender polarize, 3He cell polarizer analyze, using monitors
Thermal neutrons: 3He cell polarizer using Heussler crystals, using detectors
PASTIS setup requests a set of measurements x,y,z from ZrTi or similar

## Slide 107: Hot commissioning draft: scientific performance
Well known samples to test inelastic capability: e.g Pb or Be, Ni, Fe, MnF2, RbMnF3, single x-tals to test PSD capabilities and acquisition rates, map resolution ellipsoid
Friendly users
Troubleshooting and problem solving, challenging experiments, early success strategy
Enroll into the user programme

## Slide 108: More Questions about Work Package Specification or Hot commissioning?

## Slide 109
Thanks for your attention

## Slide 110
Detectors Group review
Plan adequate shielding of beam focusing and transmitted beam
1 deg minimum angle is challenging due to background
local instantaneous flux 1e4n/s/cm2 saturates 3He
Detection efficiency is a “should”
WPS doc says that detector tech decision has not been made
Is detector an early procurement?
Issues to be addressed afterTG2:
Kapton foils in the radial collimator
Monitors requirements in SysReq doc

## Slide 111: SAD review
Increase sample to cave ceiling distance to 3.5 - 4 m (cryomagnet + stick)
Magnetic field limits at M chopper position (detailed design, nice to know already now to set requirements of compensated magnet)
Concerns about the cave access, maybe not enough space for rolling SEE. Get lost tube and beam stop may be in the way.
PSD doc says racks are inside the cave

## Slide 112: Motion Control

## Slide 113: NOSG Review
Optics
New document ESS-0097645 to be used for calculations of substrates lifetime
Shielding
Discuss background issues during TG2
Docs do not say why the MCNP calculations are not ready
In first approximation background conditions similar to MAGIC
Out of LoS shielding could be calculated from Vitess simulations

## Slide 114: Choppers Group Review
T0 chopper outside bunker
FAN chopper
FAN chopper development
M chopper 336Hz challenging
Magnetic field limits at the M chopper
PBS and WBS and budget for provision of T0 chopper
Early procurement is strongly advised

## Slide 115: Preparation meeting for TG2 of T-REX
The signal-to-noise ratio (S/N) of the instrument is a key performance parameter for the instrument. This needs to be reflected in both the high-level scientific requirements and the high-level system requirements, where it needs to be quantified. Please prepare a proposal for presentation and discussion at the review on how you might wish to quantify the S/N and the value to state in the requirements. This is important as it will drive many of the design choices being made over the coming months.

## Slide 116: Ken Andersen Review

## Slide 117
Ei range from 2 to 160 meV
@ Ei = 3 meV
30 µeV < δE < 90 µeV
@ Ei = 100 meV
3 meV < δE < 6 meV
Intensity
≤ 109 n/s
T-REX Competitive to all existing DGCS

## Slide 118
T-REX unique with 5D mapping
instrument | E_min | E_max | ϑ_min | ϑ_max | ϕ_min | ϕ_max | L_SD | ΔΩ | max Q
AMATERAS | 1 | 80 | 5 | 110 | -16 | 23 | 4 | 1.2 | 10.2
4-SEASONS | 5 | 300 | -35 | 84 | -25 | 27 | 2.4 | 1.9 | 16.1
CNCS | 0.5 | 80 | -50 | 140 | -16 | 16 | 3.5 | 1.9 | 11.7
HYSPEC | 3.6 | 60 | 2 | 135 | -7 | 7 | 4.5 | 0.6 | 9.9
SEQUOIA | 8 | 2000 | -30 | 60 | -18 | 18 | 5.5 | 1.0 | 31.1
ARCS | 10 | 1500 | -28 | 135 | -27 | 26 | 3 | 2.6 | 49.7
MAPS | 15 | 2000 | 3 | 60 | -15 | 15 | 6 | 0.5 | 31.1
MERLIN | 7 | 2000 | -45 | 135 | -30 | 30 | 2.5 | 3.3 | 57.4
LET | 0.5 | 30 | -40 | 140 | -30 | 30 | 3.5 | 3.3 | 7.2
IN4C | 10 | 100 | 13 | 120 | -15 | 15 | 2.5 | 1.0 | 12.0
IN5 | 0.2 | 25 | -12 | 135 | -20 | 20 | 4 | 1.8 | 6.4
IN6 | 2.35 | 4.87 | 10 | 115 | -15 | 15 | 2.5 | 1.0 | 2.6
T-REX |  | 160 |  | 144 |  |  |  |  | 16.7
5D mapping
Polychromatic experiments
+ PA unique feature
Chart: E_i max (meV)
| 80.0 | 300.0 | 80.0 | 60.0 | 2000.0 | 1500.0 | 2000.0 | 2000.0 | 30.0 | 100.0 | 25.0 | 4.87
instrument | 10.1786 | 16.101 | 11.6765 | 9.9419 | 31.0646 | 49.7097 | 31.0646 | 57.3998 | 7.1503 | 12.0313 | 6.4175 | 2.5857
50% | 12.4258 | 1.3892
cost cat | 5.1378 | 0.57442
full scope | 16.7127 | 1.8685

## Slide 119
MERLIN/LET
4-SEASONS/CNCS/IN5
AMATERAS
SEQUOIA/IN4C/IN6
MAPS/
HYSPEC
Full Scope
Project scope
T-REX
T-REX Competitive solid angle
PANTHER
hot commissioning
2018

## Slide 120
 | Steel ( m ) | Concrete ( m )
< 28 m | bunker | 
28 m < L < 45 m | 0.35 | 0.3
45 m < L < 50 m | 0.25 | 0.3
50 m < L < 162 m | 0.18 m < x < 0.28 according to m-index | 0.6
162 m < L < 172 m | Experimental Cave | 
Shielding: ESS-cost calculation process
… “agreed process with ESS management that should provide a 10% error bar on the positive side meaning that the safety shielding budget should not increase by more than 10% with the available information” …
Engineering design ?
Potential savings communal shielding

## Slide 121: Expected performance
NO NEGATIVE IMPACT ON THE SCIENCE CASE
Brightness
Brightness