## Slide 1: Performance of the Multi-Grid Detector on the Time-of-Flight Spectrometer CNCS at SNS
Anton Khaplanov
On behalf of ESS Detector Group and Collaborators
www.europeanspallationsource.se
March 27, 2017

## Slide 2: Introduction
ILL:
Bruno Guerard, Jean-Claude Buffet,
Jean-Francois Clergeau, Anthony Leandri
ESS:
Anton Khaplanov, Fatima Issa, Richard Hall-Wilton, Oliver Kirstein, Tomasz Brys, Michail Anastasopoulos, Isaak Lopez Higuera, Richard Bebb, Sara Arranz, Carina Höglund*, Linda Robinson*, Susan Schmidt*
Centre for Energy Research (Hungary):
Eszter Dian
Linköping University:
Jens Birch, Lars Hultman, (also *)
SNS:
Ken Herwig, Georg Ehlers, Michelle Everett, Kevin Berry
Earlier – the participants of the CRISP project on Large-Area detectors.
Previous publications:
B4C layers:
*C. Höglund et al, J of Appl. Phys. 111, 104908 (2012)
Characterization:
*A. Khaplanov et al., arXiv:1209.0566 (2012)
*B Guerard et al., NIMA, 720, 116-121 (2013), http://dx.doi.org/10.1016/j.nima.2012.12.021iJ
*J. Correa et al., Trans. Nucl. Sc. (2013), DOI: 10.1109/TNS.2012.2227798
*A. Khaplanov et al., (2014) J. Phys.: Conf. Ser. 528 012040 doi:10.1088/1742-6596/528/1/012040
Gamma sensitivity:
*A. Khaplanov et al., JINST 8, P10025 (2013), arXiv:1306.6247
Alpha background:
*A. Khaplanov et al., JINST 10, P10019 (2015); doi:10.1088/1748-0221/10/10/P10019
Current work:
A.Khaplanov et al. “Multi-Grid Detector for Neutron Spectroscopy: Results Obtained on Time-of-Flight Spectrometer CNCS" https://arxiv.org/abs/1703.03626 (submitted to JINST)
WP 4.3: Large-Area Detectors
Horizon 2020 grant agreement 676548

## Slide 3: Outline
ESS direct geometry spectrometers
Multi-Grid detector
Multi-Grid setup at CNCS
Results from CNCS
Conclusion

## Slide 4: Time-of-Flight Spectrometers at ESS
Instrument | C-SPEC | T-REX | VOR
Sample – detector | 3.5 m | 3 m | 3 m *
Detector height | 3.5 m | 2.2 m | 3 m *
Detector area | 29 m2 | 21 m2 | 37 m2 *
3 ToF spectrometers for ESS
CSPEC and T-REX – design phase
VOR – accepted proposal
Multi-Grid as baseline detector technology
*preliminary

## Slide 5: Time-of-Flight Spectrometers
Time-of-flight information
2-5 m sample to detector distance
 Energy transfer
20-40 m2 detector area
Position resolution O(1inch)
Momentum transfer
High dynamic range
Low background
n Ei, Qi
n Ef, Qf
Incident beam
Sample
Detected ToF spectrum
Detected image

## Slide 6: CSPEC at ESS
Cold spectrometer
0.2 meV < Ei < 20 meV
29m2 detector
Horizontal coverage 5° to 135°
Vertical coverage -25° to 25°

## Slide 7: T-REX at ESS
Thermal / bispectral spectrometer
2 meV < Ei < 160 meV
21m2 detector
Horizontal coverage -36° to -1° and 1° to 144°
Vertical coverage -25° to 15°

## Slide 8: Multi-Grid Principle
B-10 film
Al substrate
n
conversion
gas
fragment 1
fragment 2
gas
n
n
Introduced at ILL, jointly developed by ILL and ESS under CRISP project, and now under BrightnESS
High total efficiency
Readout split over many layers – lower local rate
No need for resistive readout – low gain

## Slide 9: Construction of MG.CNCS in Lund
1.1m

## Slide 10: Goals of the Test on CNCS
Goals:
Test at spectrometer
Side-by-side comparison to He3
In parallel with user experiments
Dedicated tests
Long-term operation
July 2016
Feb 2017

## Slide 11: Layout of CNCS at SNS
Cold Neutron Chopper Spectrometer
Topics:
Magnetism
Structural excitations
Dynamics in confined geometries

## Slide 12: Multi-Grid installed at CNCS
Detector installed at 57° scattering angle
Inaccessible between summer and winter shutdowns
DAQ setup outside, accessible remotely
MG
He-tubes
3He 8-packs
M-G
1.1m
19cm
20cm
View from sample
2m

## Slide 13: Dedicated Measurements
Vanadium
Background
Single crystal
Fast neutrons (fission, prompt)
Water, Ice
Allowing to measure:
Energy resolution, efficiency, scattering, background sensitivity, saturation
Gamma background
(NaI detector)

## Slide 14: Time-of-Flight Spectrum
Elastic peak
Neutron energy loss
Neutron energy gain
Prompt peak
Time between source pulses

## Slide 15: Energy Reconstruction
Energy transfer:
Ef-Ei

## Slide 16: Elastic Energy Resolution

## Slide 17: Efficiency
Counts from integral of the elastic peak in dE
Efficiency measured relative to the nearest He3 tubes
Calculated tube efficiency for 6bar He3

## Slide 18: Single Crystal Reflection
Crystal
Incident beam
Reflected beam at Ei = 17.20meV
Reflected beam at Ei = 13.74meV
A single crystal reflects neutrons according to Bragg’s law:
Resulting in an intense spot seen by detector
He3
MG

## Slide 19: Single Crystal Reflection
He3
Multi-Grid
Image recorded in each detector as a function of time during elastic peak
Each frame advances 30 μs

## Slide 20: Single Crystal Reflection
He3
MG
The 6th image frame from previous slide
(Peak rate in 1 tube / 1 line of pixels ≈ 150kHz)
Position:
Charge division in He3 PSD – saturation causes loss of position resolution
Coincidence between grids and wires in Multi-Grid
Local count rate limit:
space charge in gas (*)
  x10 lower in MG per event
  x10 lower in MG due to lower gain
Dead time of electronics
depends on electronics, similar
*see presentation F. Piscitelli today 14:25

## Slide 21: Conclusions
Operation with no access for 8 months so far.
All data compared to He3. Good performance demonstrated for ToF spectrometer.
Energy resolution, efficiency, backgrounds
    shown to be comparable between MG and He3.
Better performance in high rate in Multi-Grid
JINST paper submitted (see https://arxiv.org/abs/1703.03626 for all results)
Detector currently baseline for CSPEC, T-REX at ESS
Thanks to the colleagues at SNS for this opportunity and cooperation!
Further on related simulations, see E. Dian presentation today 14:00

## Slide 23: Choice of 10B4C Layer Thickness
Efficiency optimized for λ=4Å
10B4C coating thicknesses:
7 blades with 0.5 um
7 blades with 1.0 um
3 blades with 1.5 um

## Slide 24: Aluminium Bragg Edge
Incident n
Sample
Sample environment
Detector
Scattering effects from SE and detector

## Slide 25: Gamma Background

## Slide 26: Fast Neutron Background

## Slide 27: Energy Reconstruction
Sample center (0, 0, 0)
Grid reference
(2981.9, 149.6, 1943.9)
dist
Pixel position (x, y, z)
3559.6
149.6
Ei
vi
λi
Ef
vf
λf
Reconstructed for each n
ts
td

## Slide 28: Energy Reconstruction Front Cells Only
Sample center (0, 0, 0)
Grid reference
(2981.9, 149.6, 1943.9)
dist
Pixel position (x, y, z)
149.6
Ei
vi
λi
Ef
vf
λf
ts
td
Correction for depth does not compromise data quality
Detector equivalent to a flat detector
~16% of events in 1st cells for 9Å

## Slide 29: Single Crystal Reflection

## Slide 30: Shielding Used in MG.CNCS
Aluminium
Gd2O3
B4C
PCB