## Slide 1: Characterization of Thermal Neutron Beam Monitors
Fatima Issa
www.europeanspallationsource.se
21 March 2017

## Slide 2: Outline
why Beam monitors?
Types and specifications of the used Beam monitors
results (efficiency, attenuation, gamma sensitivity,.. )
conclusion

## Slide 3: ESS instruments
Instrument Class | Instrument Sub-class
Large Scale Structures | Small Angle Scattering
 | Reflectometry
Diffraction | Powder Diffraction
 | Single crystal diffraction
Engineering | Strain scanning
 | Imaging and tomography
Spectroscopy | Direct Geometry
 | Indirect Geometry
At each instrument beam monitors will be used for diagnostics and for normalization

## Slide 4: Beam Monitor
Nuclear reactions used for thermal neutron detection
Beam monitors are thermal neutron detectors with relatively low efficiency which varies from
10-6 to 10-2 depending on the instrument requirements

## Slide 5: Types of Beam Monitor
Multi-wire proportional chamber (MWPC)
From Mirrotron
From ORDELA
MWPC Can be filled with either 3He or 14N gas

## Slide 6
Fission Chamber:
From Quantum detector
Scintillator
Types of Beam Monitor
From LND
Counting gas P10
Inner wall coated with 235U

## Slide 7: Types of Beam Monitor
Multi-wire proportional chamber (2D-MWPC)
Gas Electron Multiplier (2D-GEM)
From Mirrotron
From CDT
0.8 μm boron layer was selected for the sake of the measurement. Thinner layer will be selected for full implementation (50 -100 nm).

## Slide 8: Beam Monitors-Gamma sensitivity measurement
Beam monitor
Gamma source
neutron source
Beam monitor
MWPC from Mirrotron
Fission chamber from LND
Gamma sensitivity studied with three gamma sources with low, medium, and high energy
source | Gamma energy KeV | Gamma intensity %
57Co | 122 | 85.6
 | 136 | 10.7
133Ba | 30-35 | 117
 | 80 | 36
 | 276 | 7
 | 303 | 18
 | 356 | 62
 | 384 | 9
60Co | 1173 | 100
 | 1332 | 100

## Slide 9: Beam Monitors-Gamma sensitivity measurement
Beam monitor
Gamma source
neutron source
Beam monitor
Gamma sensitivity depends on the set threshold, fission chamber showed the lowest gamma sensitivity
source | Gamma energy KeV | Gamma intensity %
57Co | 122 | 85.6
 | 136 | 10.7
133Ba | 30-35 | 117
 | 80 | 36
 | 276 | 7
 | 303 | 18
 | 356 | 62
 | 384 | 9
60Co | 1173 | 100
 | 1332 | 100
Measuring time 10 min

## Slide 10: Beam Monitors- Neutron measurement at R2D2
Slit 1
Slit 2
Peak wavelength (Å)
400 2.00
911 0.878
711 1.12
511 1.54
822 0.943
311 2.41
933 0.804

## Slide 11: Beam Monitors- Neutron measurement
ORDELA-He
ORDELA-N
2D-GEM
2D-MWPC

## Slide 12: Beam monitors-main results
 | MWPC | MWPC | MWPC BM-100X50 | 2D-MWPC | 2D-GEM | Scintillator | Fission chamber
supplier | ORDELA | ORDELA | Mirrotron | Mirrotron | CDT | Quantum detector | LND
Isotope used for neutron capture | 3He | 14N | 3He | 3He | 10B | 6Li | 235U
Gas pressure mbar | Partial pressure 6,0795 | Partial pressure 81,06 | Partial pressure 6,5 | Partial pressure 0.4 | Total pressure 100 | ------ | Total pressure 1013,2
Filled gas | 3He+4He +CF4 | N+CF4 | 3He+CF4 | 3He+CF4 | Ar/CO2 | -------- | P10
Active Area (mm2) | 114 x 51 | 114 x 51 | 100 x 50 | 100 x 50 | Diameter 100 mm | 28 x 42 | Diameter 108.0 mm
Applied voltage (V) | 850 | 850 | 1300 | Anode at -3500V Drift at 1500V | -1000 | 650 | 300
Measured attenuation % | 4.5 | 4.4 | 2.5 | 7.3 | 10.6 | 0.4 | 3.8
Calculated attenuation % | 4.05 | 4.05 | 2.05 | 4.5 | 6.9 | 0.1 | 2.4
Measured Efficiency at 2.4Å | 1.2x10-3 | 2x10-5 | 1.1x10-3 | 1.5x10-4 | 3x10-2 | 5x10-3 | 10-4
More characterization  monitors will be customized to match ESS requirements for each instruments

## Slide 13: Conclusions
Different types of beam monitors (MWPC, 2D-GEM, scintillator, fission chamber) from different supplier have been tested
Efficiency, attenuation and gamma sensitivity were studied
Relatively high attenuation factor  mainly contributed to the window thickness
more characterization is planned (sensitivity to fast neutrons, ageing, linearity, )
monitors will be customized to match ESS requirements for each instruments

## Slide 14
Thank you
Horizon 2020 grant agreement 676548