## Slide 1
Neutron Detection in a High Rate Environment, with Reflectometry at ESS as an example
Francesco Piscitelli
on behalf of
the ESS / Wigner / Lund University / Linköping University collaboration
ICANS XXII
2017/03/27
Wigner Research Institute
Building research infrastructure and synergies for highest scientific impact on ESS
Horizon 2020 grant agreement 676548
Linköping University

## Slide 2
Counting rate definitions
4 cm

## Slide 3
Counting rate definitions
4 cm

## Slide 4
intensity
time
Counting rate definitions
4 cm
Source

## Slide 5
Counting rate definitions
4 cm
intensity
time
detector area
Flux at detector
Source

## Slide 6
intensity
time
detector area
Counting rate definitions
neutron delivery system
Wavelengths spread in time, intensities are chopped and shaped
4 cm
sample
Flux at sample
Flux at detector
Beam transportation
Source
Each instrument is different!

## Slide 7
intensity
time
detector area
Counting rate definitions
neutron delivery system
Wavelengths spread in time, intensities are chopped and shaped
4 cm
sample
Flux at sample
Flux at detector
Beam transportation
Source
Each instrument is different!

## Slide 8
Counting rate definitions
 | Global over whole detector area | Local per cm2, mm2 or pixel
Time-averaged
per s
Instantaneous
per s
*I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.

## Slide 9
Counting rate definitions
 | Global over whole detector area | Local per cm2, mm2 or pixel
intensity
Time-averaged
per s
Instantaneous
per s
time
intensity
time
intensity
time
intensity
time
*I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.

## Slide 10
Counting rate definitions
 | Global over whole detector area | Local per cm2, mm2 or pixel
intensity
Time-averaged
per s
Instantaneous
per s
time
intensity
time
intensity
time
intensity
time
*I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.
At peak is the worst condition, i.e. highest

## Slide 11
Counting rate definitions
 | Global over whole detector area | Local per cm2, mm2 or pixel
intensity
Time-averaged
per s
Instantaneous
per s
time
intensity
time
intensity
time
intensity
time
detector area
*I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.
Total number of neutrons per second on whole detector area
Relevant for bandwidth in DAQ and storage
At peak is the worst condition, i.e. highest

## Slide 12
Counting rate definitions
 | Global over whole detector area | Local per cm2, mm2 or pixel
intensity
Time-averaged
per s
Instantaneous
per s
time
intensity
time
intensity
time
intensity
time
detector area
detector area
*I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.
Total number of neutrons per second on whole detector area
Relevant for bandwidth in DAQ and storage
Total number of neutrons per second on a pixel or unit area
At peak is the worst condition, i.e. highest

## Slide 13
Counting rate definitions
 | Global over whole detector area | Local per cm2, mm2 or pixel
intensity
Time-averaged
per s
Instantaneous
per s
time
intensity
time
intensity
time
intensity
time
detector area
detector area
detector area
detector area
*I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.
Total number of neutrons per second on whole detector area
Relevant for bandwidth in DAQ and storage
Total number of neutrons per second on a pixel or unit area
At peak is the worst condition, i.e. highest
Highest instantaneous number of neutrons per second on whole detector area
Highest instantaneous number of neutrons per second in the brightest pixel or unit area
Relevant for detector technology choice

## Slide 14
Counting rate definitions
 | Global over whole detector area | Local per cm2, mm2 or pixel
intensity
Time-averaged
per s
Instantaneous
per s
time
intensity
time
intensity
time
intensity
time
detector area
detector area
detector area
detector area
*I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.
Total number of neutrons per second on whole detector area
Relevant for bandwidth in DAQ and storage
Total number of neutrons per second on a pixel or unit area
At peak is the worst condition, i.e. highest
Highest instantaneous number of neutrons per second on whole detector area
Highest instantaneous number of neutrons per second in the brightest pixel or unit area
Relevant for detector technology choice
Reflectometry
SANS
Data storage

## Slide 15
Detector requirements in terms of instantaneous counting rate capability
intensity
time
detector area
Highest instantaneous number of neutrons per second in the brightest pixel or unit area

## Slide 16
Detector requirements in terms of instantaneous counting rate capability
Direct Spectroscopy
Indirect Spectroscopy
Diffraction
Imaging
SANS
Reflectometry
1 10 100 1000
1 = state-of-the-art
Instantaneous
Rate Capability (log)
factor by which requirements
exceed state-of-the-art
Rate requirements
intensity
time
detector area
Highest instantaneous number of neutrons per second in the brightest pixel or unit area

## Slide 17
Vertical Reflectometer
Horizontal Reflectometer
Reflectometry at ESS: FREIA and ESTIA

## Slide 18
Reflectometry is a technique to study
SURFACES AND INTERFACES
Reflectometry

## Slide 19
sample
q
q
detector
I0 incoming intensity
Specular reflection
IR s. refl. intensity
Reflectometry is a technique to study
SURFACES AND INTERFACES
Reflectometry

## Slide 20
q =(4p/l) sin(q)
Neutron wavelength
Incidence angle
To measure the reflected neutrons
as a function of q
sample
q
q
detector
I0 incoming intensity
Specular reflection
IR s. refl. intensity
Reflectometry is a technique to study
SURFACES AND INTERFACES
q
Reflectometry

## Slide 21
y
x
localized
specular
reflected
intensity
(2mm x 40mm)
sample
q
q
detector
I0 incoming intensity
Specular reflection
IR s. refl. intensity
detector
Reflectometry

## Slide 22
y
x
sample
q
q
detector
I0 incoming intensity
Specular reflection
IR s. refl. intensity
detector
Reflectometry
ESTIA can work in 3 different modes:
FREIA measures 3 angles at once:
sample
localized
specular
reflected
intensity
(2mm x 40mm)

## Slide 23
Reflectometry at ESS: FREIA

## Slide 24
Reflectometry at ESS: FREIA
intensity
time
detector area
neutron delivery system
sample
Flux at sample
Flux at detector
Beam transportation
Source

## Slide 25
Reflectometry at ESS: FREIA
intensity
time
detector area
neutron delivery system
sample
Flux at sample
Flux at detector
Beam transportation
Source
q =(4p/l) sin(q)
sample
3 ranges in q measured at once without moving the sample

## Slide 26
q =(4p/l) sin(q)
Incoming Flux at sample
Reflectometry at ESS: FREIA
intensity
time
detector area
neutron delivery system
sample
Flux at sample
Flux at detector
Beam transportation
Source
sample

## Slide 27
q =(4p/l) sin(q)
Incoming Flux at sample
106
107
108 n/s/cm2
(at peak)
sample
106
107
108
106
105
104
R=100
R<10-1
R<10-4
Reflectometry at ESS: FREIA
detector
sample
3 angles

## Slide 28
q =(4p/l) sin(q)
Incoming Flux at sample
106
107
108 n/s/cm2
(at peak)
sample
106
107
108
106
105
104
R=100
R<10-1
R<10-4
Reflectometry at ESS: FREIA
detector
+ divergence
+ footprint on detector
to get the rate per mm2
sample
3 angles

## Slide 29
Detector requirements for Reflectometry at ESS

## Slide 30
Detector requirements for Reflectometry at ESS
3 times
state-of-the-art detectors
10-100 times
state-of-the-art detectors

## Slide 31
4 cm
What can we do to meet the requirments?

## Slide 32
neutron
converter
10B, 3He, 6Li, …
charged particles
detection
electric signals
electronics and
DAQ
Neutron detector schematic

## Slide 33
neutron
converter
10B, 3He, 6Li, …
charged particles
detection
electric signals
electronics and
DAQ
Neutron detector schematic
… rate limitation mainly comes from here!
Not really important which converter you use but how you optimize the detection process

## Slide 34
neutron
converter
10B, 3He, 6Li, …
charged particles
detection
electric signals
electronics and
DAQ
Neutron detector schematic
… rate limitation mainly comes from here!
Not really important which converter you use but how you optimize the detection process
gain ~ charge to evacuate
rate capability

## Slide 35
Publications about the Multi-Grid:
J. Birch et al., 10B4C Multi-Grid as an alternative to 3He for large area neutorn detectors, IEEE TNS, vol. 60, no. 2, 2013.
J. Birch et al., In-beam test of the Boron-10 Multi-Grid
neutron detector at the IN6 time-of-flight spectrometer
at the ILL, JoP, vol. 528, 012040, 2014.
A. Khaplanov et al., Multi-Grid Detector for Neutron Spectroscopy: Results Obtained on Time-of-Flight Spectrometer CNCS, arXiv: 1703.03626, 2017
(submitted to JINST)
neutron
10B4C layer
substrate
10B4C layer
gas
volume
gas
volume
2. Grazing angle (<10˚)
substrate
neutron
10B-layer-based detectors
Publications about the Multi-Blade:
F. Piscitelli et al., Study of a high spatial resolution
10B-based thermal neutron detector for application
in neutron reflectometry: the Multi-Blade prototype,
JINST vol. 9, P03007, 2014
F. Piscitelli et al., The Multi-Blade Boron-10-based
Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST vol.12, P03013, 2017.
1. Multi-layer

## Slide 36
Readout options for gaseous detectors
neutron
charged particles
detection
electric signals
electronics and
DAQ
GEM
Micromegas
Microstrips
MWPC
converter
10B, 3He, 6Li, …

## Slide 37
Readout options for gaseous detectors
neutron
charged particles
detection
electric signals
electronics and
DAQ
GEM
Micromegas
Microstrips
MWPC
converter
10B, 3He, 6Li, …
Counting rate capability limited mainly by space charge effect and double hit time resolution

## Slide 38
A. Ravazzani et al. , “Characterisation of proportional counters,” Radiation Measurements , vol. 41, no. 5, pp. 582 – 593, 2006.
A. Andronic et al. , “A comprehensive study of rate capability in multi-wire proportional chambers,” Journal of Instrumentation , vol. 4, no. 10, p. P10014, 2009.
H. Sipilä and V. Vanha-Honko, “The counting rate dependent gain shift and the space charge in proportional counters,” Nuclear Instruments and Methods , vol. 153, no. 2, pp. 461 – 464, 1978.
E. Mathieson, “Dependence of gain on count rate, due to space charge, in coaxial and multiwire proportional chambers,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 249, no. 2,
pp. 413 – 420, 1986.
E. Mathieson and G. Smith, “Gain reduction due to space charge in a multiwire proportional chamber irradiated by a uniform beam of rectangular section,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 316, no. 2, pp. 246 – 251, 1992.
Y. Ivaniouchenkov et al. , “Breakdown limit studies in high-rate gaseous detectors,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 422, no. 1–3, pp. 300 – 304, 1999.
S. Koperny and T. Z. Kowalski, “11th topical seminar on innovative particle and radiation detectors (IPRD08) performance of proportional counters under high count rate, high gas gain and at high working gas pressure,” Nuclear Physics B - Proceedings Supplements , vol. 197, no. 1, pp. 370 – 373, 2009.
P. Fonte, “Which gaseous detector is the best at high rates?,” SLAC-JOURNAL-ICFA-16-2 .
E. P. Cippo et al. , “A GEM-based thermal neutron detector for high counting rate
applications,” Journal of Instrumentation , vol. 10, no. 10, p. P10003, 2015.
GEM
Micromegas
Microstrips
MWPC
Readout options for gaseous detectors

## Slide 39
A. Ravazzani et al. , “Characterisation of proportional counters,” Radiation Measurements , vol. 41, no. 5, pp. 582 – 593, 2006.
A. Andronic et al. , “A comprehensive study of rate capability in multi-wire proportional chambers,” Journal of Instrumentation , vol. 4, no. 10, p. P10014, 2009.
H. Sipilä and V. Vanha-Honko, “The counting rate dependent gain shift and the space charge in proportional counters,” Nuclear Instruments and Methods , vol. 153, no. 2, pp. 461 – 464, 1978.
E. Mathieson, “Dependence of gain on count rate, due to space charge, in coaxial and multiwire proportional chambers,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 249, no. 2,
pp. 413 – 420, 1986.
E. Mathieson and G. Smith, “Gain reduction due to space charge in a multiwire proportional chamber irradiated by a uniform beam of rectangular section,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 316, no. 2, pp. 246 – 251, 1992.
Y. Ivaniouchenkov et al. , “Breakdown limit studies in high-rate gaseous detectors,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 422, no. 1–3, pp. 300 – 304, 1999.
S. Koperny and T. Z. Kowalski, “11th topical seminar on innovative particle and radiation detectors (IPRD08) performance of proportional counters under high count rate, high gas gain and at high working gas pressure,” Nuclear Physics B - Proceedings Supplements , vol. 197, no. 1, pp. 370 – 373, 2009.
P. Fonte, “Which gaseous detector is the best at high rates?,” SLAC-JOURNAL-ICFA-16-2 .
E. P. Cippo et al. , “A GEM-based thermal neutron detector for high counting rate
applications,” Journal of Instrumentation , vol. 10, no. 10, p. P10003, 2015.
GEM
Micromegas
Microstrips
MWPC
Readout options for gaseous detectors
Most of them studied with
X-rays but not with neutrons

## Slide 40
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
Readout options for gaseous detectors
Rate
Gain
(or total charge in the detector)
Ideal case
Rate performance

## Slide 41
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
Readout options for gaseous detectors
Rate
Ideal case
reality
Gain drops
Gain
(or total charge in the detector)
Rate performance

## Slide 42
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
Readout options for gaseous detectors
Rate
Ideal case
Higher gain
Lower gain
Gain
(or total charge in the detector)
Rate performance

## Slide 43
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
Readout options for gaseous detectors
Rate
Ideal case
Higher gain
Lower gain
Gain
(or total charge in the detector)
gain ~ charge to evacuate
rate capability
Rate performance

## Slide 44
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC

## Slide 45
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC

## Slide 46
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC
Does not depend on the 3He but to the usually employed charge division readout that needs charge to get the position resolution

## Slide 47
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors
neutron
alpha
1470 keV
< 10 fC
G = 15
converter
10B
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC
Does not depend on the 3He but to the usually employed charge division readout that needs charge to get the position resolution

## Slide 48
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors
neutron
alpha
1470 keV
< 10 fC
G = 15
converter
10B
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC
MWPC
MM
Does not depend on the 3He but to the usually employed charge division readout that needs charge to get the position resolution
G
MSGC
studied with X-rays

## Slide 49
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors
neutron
alpha
1470 keV
< 10 fC
G = 15
converter
10B
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC
MWPC
MM
MSGC
equivalent neutron
charge event
for 10B readout
equivalent neutron
charge event
for 3He readout
Does not depend on the 3He but to the usually employed charge division readout that needs charge to get the position resolution
G
studied with X-rays
GEM

## Slide 50
GEM
Micromegas
(MM)
Microstrips (MSGC)
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors
neutron
alpha
1470 keV
< 10 fC
G = 15
converter
10B
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC
MWPC
MM
equivalent neutron
charge event
for 10B readout
MWPC can reach >104 Hz/mm2
equivalent neutron
charge event
for 3He readout
Does not depend on the 3He but to the usually employed charge division readout that needs charge to get the position resolution
G
MSGC
studied with X-rays
GEM

## Slide 51
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors: MWPC
neutron
alpha
1470 keV
< 10 fC
G = 15
converter
10B
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC
equivalent neutron
charge event
for 10B readout
equivalent neutron
charge event
for 3He readout
G
100 1000 10000 100000
Rate (Hz/mm2)

## Slide 52
MWPC
X-ray
6 keV
< 0.05 fC
(in Ar/CO2)
Readout options for gaseous detectors: MWPC
neutron
alpha
1470 keV
< 10 fC
G = 15
converter
10B
neutron
proton + t
770 keV
< 3 fC
G = 500
converter3He
G = 3000
G = 30000
Final charge =
0.15pC
1.5pC
equivalent neutron
charge event
for 10B readout
equivalent neutron
charge event
for 3He readout
G
100 1000 10000 100000
Rate (Hz/mm2)
gain ~ charge to evacuate
rate capability

## Slide 53
The Multi-Blade project
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 54
The Multi-Blade project
Wigner Research Institute
Lund
Budapest
The key objective of WP4 is the technological
evolution of neutron detectors in terms of
resolution, intensity and dimensions.
3 years
Task 4.2 Neutron Detectors – The Intensity Frontier
Linköping University
Linköping

## Slide 55
neutrons
5 degrees
θ
Efficiency 45% at 2.5Å
A single Boron layer inclined at 5 degrees
Efficiency
substrate
10B4C
The Multi-Blade project
q = 5 degrees

## Slide 56
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
The Multi-Blade project
MWPC
q = 5 degrees

## Slide 57
Why the counting rate capability is improved?
The intensity is spread over a wider surface (5 degrees ~ factor x10)
    Thin gap MWPC (4mm)
      Low gas gain operation G~20
(max 0.2pC avalanches)
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
The Multi-Blade project
MWPC
q = 5 degrees

## Slide 58
Why the counting rate capability is improved?
The intensity is spread over a wider surface (5 degrees ~ factor x10)
    Thin gap MWPC (4mm)
      Low gas gain operation G~20
(max 0.2pC avalanches)
wires X
substrate
neutrons
θ
strips Y
10B4C
The Multi-Blade project
q = 5 degrees
rate
ESS requirement
Rate (Hz/mm2)
100 1000 10000 100000

## Slide 59
Why the counting rate capability is improved?
The intensity is spread over a wider surface (5 degrees ~ factor x10)
    Thin gap MWPC (4mm)
      Low gas gain operation G~20
(max 0.2pC avalanches)
wires X
substrate
neutrons
θ
strips Y
10B4C
The Multi-Blade project
q = 5 degrees
rate
ESS requirement
gain ~ charge to evacuate
rate capability
Rate (Hz/mm2)
100 1000 10000 100000
equivalent neutron
charge event
for 10B readout

## Slide 60
Why the counting rate capability is improved?
The intensity is spread over a wider surface (5 degrees ~ factor x10)
    Thin gap MWPC (4mm)
      Low gas gain operation G~20
(max 0.2pC avalanches)
wires X
substrate
neutrons
θ
strips Y
10B4C
The Multi-Blade project
q = 5 degrees
rate
ESS requirement
Rate (Hz/mm2)
100 1000 10000 100000
equivalent neutron
charge event
for 10B readout
gain ~ charge to evacuate
rate capability

## Slide 61
neutrons
Old Design
cassette
10B4C
substrate
Materials evaluation
Reduce the scattering
Inclination makes a factor 10 thicker layers!
NO MPGD

## Slide 62
neutrons
Old Design
cassette
strips Y
wires X
10B4C
substrate
Materials evaluation
Reduce the scattering
Inclination makes a factor 10 thicker layers!
30um kapton is 300um at 5deg
MWPC

## Slide 63
neutrons
neutrons
strips Y
wires X
Old Design New Design
cassette
cassette
substrate
10B4C
strips Y
wires X
10B4C
substrate
Materials evaluation
Reduce the scattering
q = 5 degrees
θ
MWPC
MWPC

## Slide 64
neutrons
neutrons
strips Y
wires X
Old Design New Design
cassette
cassette
substrate
10B4C
strips Y
wires X
10B4C
substrate
Materials evaluation
Reduce the scattering
X
2.5Å
99% absorption with 7μm
θ
q = 5 degrees
MWPC
MWPC

## Slide 65
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
The Multi-Blade project
MWPC
q = 5 degrees

## Slide 66
Mechanical design
amplifier board
amplifier board
cathode strips
anode wires
holding frame
10B4C substrate
neutrons
100x140mm2 active area demonstrator

## Slide 67
amplifier board
amplifier board
cathode strips
anode wires
holding frame
10B4C substrate
10B4C substrate
holding frame
cathode strips
cathode strips
The Multi-Blade detector

## Slide 68
10B4C substrate
holding frame
cathode strips
cathode strips
The Multi-Blade detector

## Slide 69
The Multi-Blade detector

## Slide 70
140mm
100mm
The Multi-Blade detector
neutrons
neutrons
Multi-Blade

## Slide 71
Results
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 72
Efficiency
64% @ 5.1Å
measured
Calculated*
58% @ 4.2Å
45% @ 2.5Å
*F. Piscitelli and P. van Esch,
Analytical modeling of thin film neutron converters and its application to thermal neutron gas detectors
J. Instrum. 8, P04020 (2013)
arXiv:1302.3153
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 73
PHS and Gain
Measured @ 4.1Å 5 degrees
measured
calculated*
*F. Piscitelli and P. van Esch,
Analytical modeling of thin film neutron converters and its application to thermal neutron gas detectors
J. Instrum. 8, P04020 (2013)
arXiv:1302.3153
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).
Gas Gain ~ 20 (max 0.2pC avalanche charge)
(contributes to improve the counting rate capability)

## Slide 74
PHS and Gain
Gas Gain ~ 20 (max 0.2pC avalanche charge)
(contributes to improve the counting rate capability)
Measured @ 4.1Å 5 degrees
measured
calculated*
*F. Piscitelli and P. van Esch,
Analytical modeling of thin film neutron converters and its application to thermal neutron gas detectors
J. Instrum. 8, P04020 (2013)
arXiv:1302.3153
>105 Hz/mm2 is achievable
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).
100 1000 10000 100000
Rate (Hz/mm2)
ESS requirement
equivalent neutron
charge event
for 10B readout

## Slide 75
Spot
Counting rate capability
No saturation observed!
Slit
Rate per mm2
Rate per channel
17kHz/ch
1.6kHz/mm2
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 76
24mm
Images
14mm
24mm
14mm
Raw image from the detector
Mask
Spatial Resolution FWHM ~ 0.5mm
Spatial Resolution FWHM ~ 2.5mm
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 77
Images
60mm
35mm
60mm
35mm
Raw image from the detector (log scale)
Mask
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 78
Summary

## Slide 79
Summary

## Slide 80
Summary

## Slide 81
Summary
Matching ESS requirements
Focus on next steps

## Slide 82
Summary
Matching ESS requirements
x3 better than
state-of-the-art
x10 better than
state-of-the-art

## Slide 83
Summary
Matching ESS requirements
x3 better than
state-of-the-art
x10 better than
state-of-the-art
A new detector is going to be tested this year on a Reflectometer

## Slide 84
Thank you.
Wigner Research Institute
Linköping University
Building research infrastructure and synergies for highest scientific impact on ESS
Horizon 2020 grant agreement 676548

## Slide 85
Backup slides

## Slide 86
References:
I. Stefanescu et al. ,“Neutron detectors for the ESS diffractometers,” Journal of Instrumentation , vol. 12, no. 01, p. P01019, 2017.
O. Kirstein et al. , “Neutron position sensitive detectors for the ESS,” arXiv:1411.6194 ,vol. Proceedings of the 23rd International Workshop on Vertex
Detectors, 15-19 September 2014, Macha Lake, The Czech Republic, no. PoS(Vertex2014)029, 2014.
H. Wacklin. FREIA: Reflectometer concept for fast kinetics at ESS - ESS instrument proposal –
https://europeanspallationsource.se/sites/default/files/freia_proposal.pdf, 2014.
J. Stah, U. Filges and T. Panzner, “Focusing specular neutron reflectometry for small samples,” Eur. Phys. J. Appl. Phys. , vol. 58, no. 1, p. 11001, 2012.
A. Ravazzani et al. , “Characterisation of proportional counters,” Radiation Measurements , vol. 41, no. 5, pp. 582 – 593, 2006.
B. Guerard, R. Hall-Wilton, and F. Murtas, “Prospects in MPGDs development for neutron detection,” 2014 - arXiv:1410.0107.
C. Höglund et al. , “B4C thin films for neutron detection,” Journal of Applied Physics ,vol. 111, no. 10, 2012.
A. Andronic et al. , “A comprehensive study of rate capability in multi-wire proportional chambers,” Journal of Instrumentation , vol. 4, no. 10, p. P10014, 2009.
H. Sipilä and V. Vanha-Honko, “The counting rate dependent gain shift and the space charge
in proportional counters,” Nuclear Instruments and Methods , vol. 153, no. 2, pp. 461 – 464,
1978.
E. Mathieson, “Dependence of gain on count rate, due to space charge, in coaxial and multiwire proportional chambers,” Nuclear Instruments and Methods in Physics Research
Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 249, no. 2, pp. 413 – 420, 1986.
E. Mathieson and G. Smith, “Gain reduction due to space charge in a multiwire proportional chamber irradiated by a uniform beam of rectangular section,” Nuclear Instruments
And Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment , vol. 316, no. 2, pp. 246 – 251, 1992.
Y. Ivaniouchenkov et al. , “Breakdown limit studies in high-rate gaseous detectors,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators,
Spectrometers, Detectors and Associated Equipment , vol. 422, no. 1–3, pp. 300 – 304, 1999.
S. Koperny and T. Z. Kowalski, “11th topical seminar on innovative particle and radiation detectors (IPRD08) performance of proportional counters under high count rate,
high gas gain and at high working gas pressure,” Nuclear Physics B - Proceedings Supplements , vol. 197, no. 1, pp. 370 – 373, 2009.
P. Fonte, “Which gaseous detector is the best at high rates?,” SLAC-JOURNAL-ICFA-16-2 .
M. Titov, “Perspectives of micro-pattern gaseous detector technolologies for future physics projects,” 2013 - arXiv:1308.3047.
E. P. Cippo et al. , “A GEM-based thermal neutron detector for high counting rate applications,” Journal of Instrumentation , vol. 10, no. 10, p. P10003, 2015.

## Slide 87
Al Ti SS
Planarity is an issue on large surfaces on Al
MB-16
Ti

## Slide 88
~7 mm single-side
1mm 2mm 3mm
Planarity is an issue on large surfaces
Materials evaluation
10B4C
Si
10B4C good adhesion
1mm
2mm
100x100mm2
10B4C
Al2O3
Al2O3 substrate
Al substrate