## Slide 1: Multi-Grid
Anton Khaplanov

## Slide 2: CSPEC at ESS
Cold spectrometer
0.2 meV < Ei < 20 meV
29m2 detector
Horizontal coverage 5° to 135°
Vertical coverage -25° to 25°

## Slide 3: T-REX at ESS
Thermal / bispectral spectrometer
2 meV < Ei < 160 meV
21m2 detector
Horizontal coverage -36° to -1° and 1° to 144°
Vertical coverage -25° to 15°

## Slide 4: Application of Multi-Grid: Time-of-Flight Spectrometers
Time-of-flight information
2-5 m sample to detector distance
 Energy transfer
20-40 m2 detector area
Position resolution O(1inch)
Momentum transfer
High dynamic range
Low background
n Ei, Qi
n Ef, Qf
Incident beam
Sample
Detected ToF spectrum
Detected image
IN5 at ILL

## Slide 5: Multi-Grid: How it Works
n
Low cost for large area
Modular
High total efficiency
Many layers – lower local rate
No need for resistive readout – low gain
n
n
Introduced at ILL, jointly developed by ILL and ESS under CRISP project,
now under BrightnESS

## Slide 6: Multi-Grid Concept for Data Analysis
Voxels in 3D
List mode data: [time, x, y, z]

## Slide 7: Multi-Grid installed at CNCS
Detector installed at 57° scattering angle
Inaccessible between summer and winter shutdowns
DAQ setup outside, accessible remotely
MG
He-tubes
3He 8-packs
M-G
1.1m
19cm
20cm
View from sample
2m

## Slide 8: Energy Reconstruction
Energy transfer:
Ef-Ei

## Slide 9: Some Results (Optional)

## Slide 10: Time-of-Flight Spectrum
Elastic peak
Neutron energy loss
Neutron energy gain
Prompt peak
Time between source pulses

## Slide 11: Methods Used for Tests
Vanadium
Counts over time, background
Single crystal
High local rate
Fast neutrons (fission, prompt)
Water, Ice
Allowing to measure:
Energy resolution – efficiency, scattering – background sensitivity – saturation
Gamma background
(NaI detector)
x2 lower fast n counts in MG
Rate and background analysis

## Slide 12: Elastic Energy Resolution

## Slide 13: Efficiency
Counts from integral of the elastic peak in dE
Efficiency measured relative to the nearest He3 tubes
Error bars: 5-8%, agrees with prediction

## Slide 14: Single Crystal Reflection
Crystal
Incident beam
Reflected beam at Ei = 17.20meV
Reflected beam at Ei = 13.74meV
A single crystal reflects neutrons according to Bragg’s law:
Resulting in an intense spot seen by detector
He3
MG

## Slide 15: Single Crystal Reflection
He3
Multi-Grid
Image recorded in each detector as a function of time during elastic peak
Each frame advances 30 μs
No loss of position resolution or saturation observed in Multi-Grid

## Slide 17: MG Relative CNCS He3 detectors

## Slide 18: Q Cut Comparisons for Water Sample
Analysis by Wiebke Lohstroh

## Slide 19: What is still not done for MG.CNCS data analysis
~DONE:
Automated pixel assignment
Pixel encoding gain shifts corrections
NOT DONE:
MG visualization alongside He3
dE, dQ calculation, comparison to He3 for arbitrary run
NOT DONE, but tools exist (by Morten / Edgars):
Statistical analysis over time under constant conditions
Search for “interesting” runs

## Slide 20: What is coming next: MG Thermal Instrument Test

## Slide 21: Energy Ranges for ESS Spectrometers
Energy ranges of the first ESS ToF spectrometers
Cold range, 0.76 to 15 meV measured
Measurements needed for thermal range, up to 160 meV
 | CSPEC | T-REX
Typical initial λ, Å (meV) | 2 to 15 Å (20 to 0.36 meV) | 0.7 to 6.4 Å (160 to 2 meV)
Incident energies used at CNCS
(M. Stone et al.)

## Slide 22: Efficiency Optimizations
Optimized grid
1.5μm
0.5μm
1.0μm
n
Layer thicknesses in MG.CNCS (16 blades):
7 blades 0.5μm,
7 blades 1.0μm,
3 blades 1.5μm
Proposed for a thermal detector (20 blades):
4 blades 1.0μm,
10 blades 1.25μm,
6 blades 2.0μm.
Efficiency as a function of wavelength for MG optimized for cold vs. thermal spectrum
Cold – optimization centered on 4Å
Thermal – optimization centered on 1Å
Would like to test:
Would like to test:
Tested at MG.CNCS

## Slide 23: MG Test on Thermal Instrument

## Slide 24: MG Test on Thermal Instrument

## Slide 25: MG Test on Thermal Instrument
Electronics plan:
New front-end from Mesytec
No longer using analogue channel encoding
Will provide (likely):
[module #,
channel #1, amplitude #1,
channel #2, amplitude #2,
…,
time]
In each event: every channel over threshold
Some channels are grids, some are wires
Each wire + grid combination corresponds to a unique 3D coordinate
1 of these packs for each module
(9 modules total)

## Slide 26: MG.SEQ schedule (DAQ-related)
Small-scale electronics test – by beginning of Nov 2017
event format known
sample data available
Full set of electronics available – by beginning of Feb 2018
MG.SEQ beam / source / background – March 2018 *
data from full detector available
Installation at SEQUOIA – April 2018 *
Installation coordinates available
First beam data – May 2018 and onward *
beam data
He3 data
*(fall back by ~2 month possible)

## Slide 27: MG.SEQ DAQ Wishes
Continually processing incoming data
Should be possible to see:
current rate
current image (3D, 1-2D projections)
current ToF spectrum
For a certain time index range:
rate, image, ToF, dE, dQ, …
dE, dQ will require information on Ei.
Can this be obtained automatically from instrument?
>1 algorithm for reconstructing position from grid #:
first channel
highest amplitude channel
Center of Gravity, ....
Plotting side-by-side with He3 data
Can this be obtained automatically from instrument?

## Slide 28: MG Test on Thermal Instrument