LWU BPMs signal analysis and requirements European Spallation Source ESS ERIC

September 30th, 2016

| Date | Revision | Description | Authors |
| --- | --- | --- | --- |
| 01 – July - 2016 | 0.1 | Initial draft. | Rafael A. Baron, |
| 03 – August - 2016 |  | Included mechanical discussions | Rafael A. Baron, |

Table of Contents

1	General information	4

2	BPM Specifications	4

Refer to document [1].	4

3	LWU BPM schedule	4

3.1	Delivery planning	5

3.2	Project Quality Plan	6

Refer to the documentation [1].	6

4	LWU BPM Mechanical design	6

4.1	BPM housing for the 100 mm chamber	6

4.2	BPM housing for the 57 mm chamber	7

5	LWU BPM Electromagnetic and thermal design	10

5.1	Electromagnetic performance analysis	10

5.2	Simulation Design BPM Type B	10

a)	Simulation Design BPM Type C	12

5.3	Power dissipation analysis	14

5.4	BPM signal simulations for multi-bunch operation	15

5.4.1	BPM signal analysis - Case A	15

5.4.2	BPM signal analysis - Case B	17

5.4.3	BPM signal analysis - Case C	19

5.4.4	BPM signal analysis - Case D	21

5.5	BPM electronics performance	23

8	References	24

# General information

This document is intended to describe the ESS LWU button BPM design in terms of vacuum, mechanics and system performance. LWUs BPMs specifications, schedule, vacuum design, mechanical design and BPM system performance are presented on this documentation. Discussions about the designs for Spokes, MBL, HBL, UHB, A2T, DogLg and DumpL are presented and discussed.

# BPM Specifications

# Refer to document [1].

# LWU BPM schedule

LWU BPM feedthroughs schedule requested by Daresbury

| Task Name | BPM Type | Finish | status |
| --- | --- | --- | --- |
| 20 Pickups |  | Wed 31/08/16 | Delivery – mid September |
| 20 Pickups |  | Fri 30/09/16 |  |
| 20 Pickups |  | Mon 31/10/16 |  |
| 20 Pickups |  | Wed 30/11/16 |  |
| 20 Pickups |  | Fri 30/12/16 |  |
| 20 Pickups |  | Tue 31/01/17 |  |
| 20 Pickups |  | Tue 28/02/17 |  |
| 20 Pickups |  | Fri 31/03/17 |  |
| 20 Pickups |  | Fri 28/04/17 |  |
| 20 Pickups |  | Wed 31/05/17 |  |
| 20 Pickups |  | Fri 30/06/17 |  |
| 20 Pickups |  | Mon 31/07/17 |  |
| 20 Pickups |  | Thu 31/08/17 |  |
| 20 Pickups |  | Fri 29/09/17 |  |
| 20 Pickups |  | Tue 31/10/17 |  |
| 12 Pickups |  | Thu 30/11/17 |  |

| Description | Spoke | MBL | HBL | UHB | Dogleg | A2T | DumpL | Qty BPMs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Type A (100 mm ID, 108 mm OD chamber) |  | 9 | 21 | 16 | 6 | 6 | 6 | 64 |
| Type B (57.1 mm ID, 63.5 mm OD chamber) | 14 |  |  |  |  |  |  | 14 |

Mitigations about how to avoid the delay is being done together with Daresbury and DESY. Spokes, MBL, HBL, UHB, A2T,

## Delivery planning

BPM feedthroughs manufactured by BC-Tech Mitigations about how to avoid the delay is being done together with Daresbury and DESY.

A contract for the feedthroughs production is under revision at DESY for approval.

The deliverables by DESY are described in the production contract and it is described on the following table.

DESY shall deliver the following in accordance with the table below:

| Phase | Deliverable | Destination |
| --- | --- | --- |
| Project Design Phase | The design and simulation files for the two BPM models specified by ESS | ESS |
| Project Production Phase | Production of 376 feedthroughs: consisting of 308 of type C and 68 of type B | Daresbury |
| Project Production Phase | Production documentation including ultra-high vacuum (“UHV”) and radio frequency (“RF”) test results of all the BPM feedthroughs units | ESS |
| Project Production Phase | Weekly manufacturing status reports | ESS |
| Project Production Phase | Documßentation regarding the feedthrough manufacturing process for each serial number | ESS |
| Project Production Phase | Documentation regarding select pairs of feedthroughs with similar RF characteristics | ESS |

Table: Deliverables from DESY for the LWU feedthroughs project.

The schedule described in the production contract with DESY is the following, which is in accordance to the LWU assembly schedule:

| Deliverable | Deadline |
| --- | --- |
| 52 Type C feedthroughs | 2017/01/31 |
| 256 Type C feedthroughs | 2017/03/31 |
| 68 Type B feedthroughs | 2017/12/31 |

Table: Delivery dates from DESY for the LWU feedthroughs.

Figure: Milestones for the LWU feedthroughs production

## Project Quality Plan

# Refer to the documentation [1].

# LWU BPM Mechanical design

Two types of buttons have been designed for the LINAC Warm Units in order to satisfy the beam physics specifications of position resolution and phase. For the LWUs, the specifications are divided into many different machine sections and they should be separately provided: Spokes, MBL, HBL, UHB, A2T, DogLg and DumpL.

## BPM housing for the 100 mm chamber

## BPM housing for the 57 mm chamber

Figure: BPM button type xxx.

# LWU BPM Electromagnetic and thermal design

## Electromagnetic performance analysis

The electromagnetic performance of the LWU button BPMs were analysed in order to identify the voltage level expected for each beam operation mode and to identify potential issues related to signal quality and heating due to the trapped modes around the BPM feedthrough geometry.

As described, two different types of BPMs were designed to satisfy the BPM system specifications. The simulation analyses for both cases are presented.

## Simulation Design BPM Type B

The LWUs will be designed considering two different chamber diameters. The chamber intended to be used on the has 57.1 mm diameter, and the BPM button design was made to provide enough signal to the BPM electronics in order to satisfy the Beam Physics requirements for all the beam operation modes.

Figure: Model of the BPM sensor for the 57.1 mm chamber.

Figure: Model for the electromagnetic simulation model of the BPM sensor for the 57.1 mm chamber.

Figure: Frequency domain simulations for a 2.4 mm bunch length, centered beam. In red, the signal response of the new button design, and in the green the signal response for the old button design.

Figure: Frequency domain simulations for a 2.4 mm bunch length, 15 mm off-centered beam. The sensitivity factor of the BPM is 13.63 mm, or 1.277 dB / mm of beam displacement. In red, the signal response of the new button design, and in the green the signal response for the old button design.

Figure: Frequency domain simulations for a 2.1 mm bunch length, centered beam. In red, the signal response of the new button design, and in the green the signal response for the old button design.

Figure: Frequency domain simulations for a 2.1 mm bunch length, 15 mm off-centered beam. The sensitivity factor of the BPM is 14.49 mm, or 1.201 dB / mm of beam displacement. In red, the signal response of the new button design, and in the green the signal response for the old button design.

- Simulation Design BPM Type C

Figure: Model of the BPM sensor for the 100 mm chamber.

Figure: Model for the electromagnetic simulation model of the BPM sensor for the 100 mm chamber.

Figure: Frequency domain simulations for a 2.1 mm bunch length, 26 mm off-centered beam. The sensitivity factor of the BPM is 25.01 mm, or 0.695 dB / mm of beam displacement. In red, the signal response of the new button design, and in the green the signal response for the old button design.

Figure: Frequency domain simulations for a 1.1 mm bunch length, centered beam. In red, the signal response of the new button design, and in the green the signal response for the old button design.

Figure: Frequency domain simulations for a 1.1 mm bunch length, 26 mm off-centered beam. The sensitivity factor of the BPM is 25.90 mm, or 0.671 dB / mm of beam displacement. In red, the signal response of the new button design and in the green the signal response for the old button design.

## Power dissipation analysis

The power dissipation analysis is made based on the ssss analysis, which considers the electromagnetic power stored at the geometries of the button which traps the high frequency components of the beam that interacts with the BPM housing structure. This effect is very critical for circular machines, where there is not enough time for the trapped modes to be dissipated on the walls of the mechanical structure.

For the ESS LWU Beam position monitor, the most critical condition, as can be seen in the previous section, is for the larger button, where the higher order modes are at relatively small frequencies due to the large size of the button.

The button design is not optimal for the application however the schedule constraints limited the time and resources available for the evaluation of an optimum BPM sensor for the ESS LINAC. Another BPM types can be proposed such as low frequency cavity BPMs, which can trap the modes.

An analysis of the dissipated power at the feedthrough for the 217 MeV beam energy was evaluated and the results are shown in the following figure. For this specific simulation, a beam with x and y coordinates were used.

Figure: Thermal simulations showing the dissipated power per pulse on the BPM button.

The simulated quality factor for the 1.75 GHz coupled mode is 0.8333, resulting in a decay time of 151.6 ps. Considering the bunch spacing of 2.84 ns, the mode already decayed when the arrival of the next bunch. The calculated average dissipated power on the feedthrough is around 5 mW.

Figure: Dissipated power per pulse as function of the off-centered beam, taking the mechanical center as reference.

The dissipated power can be estimated by considering the repetition rate of the ESS LINAC pulse, pulse length, and the repetition frequency.

P=fpulse.tpulse_length.fRF.kloss.Q2

Where,

- P is the average dissipated power

- fpulse is the ESS LINAC pulse repetition rate, 14 Hz

- tpulse_length is the pulse length, nominal 2.86 ms

- fRF is the cavity RF frequency, 352 MHz

- kloss is the loss factor, simulated as function of the beam position

- Q is the simulated quality factor of the cavity that traps the mode

The calculated average dissipated power due to the 1.75 GHz trapped mode is then:

P=5 mW

This average dissipated power is considered negligible since the temperature increase is in the order of tenths of a degree Celsius.

## BPM signal simulations for multi-bunch operation

The button BPMs will be used in different sections of the LINAC Warm Units (LWUs). Some of the beam characteristics changes for different parts of the machine, and they are considered by analysing 4 cases, where bunch length and beam energies varies. These cases are analysed for single bunch (electromagnetic numerical simulations) and multi bunch analysis (analytical simulations).

| Case | Chamber diameter (mm) | Button diameter (mm) | Beam energy (MeV) |
| --- | --- | --- | --- |
| A | 57,1 | 28 | 92 |
| B | 57,1 | 28 | 217 |
| C | 100 | 46 | 217 |
| D | 100 | 46 | 2000 |

General information about the analytical and electromagnetic simulations:

- Total charge: Q=62.5 mA / 352 MHz ~ 178 pC.

- 5 us of beam pulse is the minimum pulse width considered for this analysis.

- The analytical simulations consider the signal model presented on section 1.

- The cable length was considered to be 60 meters, and the cable frequency response is similar to.

### BPM signal analysis - Case A

Simulation parameters:

| Case | Chamber diameter (mm) | Button diameter (mm) | Beam energy (MeV) |
| --- | --- | --- | --- |
| A | 57,1 | 28 | 92 |

  - Electromagnetic numerical simulation

Figure: Electromagnetic simulations for single pass beam. Time and frequency domain signals for single pass are presented.

  - Analytical analysis for multi-bunch operation

Figure: Analytical simulation for case A, where time and frequency domain simulations are presented. The time domain simulations can be compared to the electromagnetic simulations presented on 1-a. The frequency domain simulations are for multi-bunch beam, repeated at a frequency of 352 MHz.

Figure: Signal power available at the BPM button and BPM electronics as function of the beam current for 5 us pulse, with a 352 MHz bunch repetition rate. The simulation on the left side shows the signal power available in the BPM button and after cables, in dBm.

### BPM signal analysis - Case B

Simulation parameters:

| Case | Chamber diameter (mm) | Button diameter (mm) | Beam energy (MeV) |
| --- | --- | --- | --- |
| B | 57,1 | 28 | 217 |

- Electromagnetic numerical simulation

Figure: Electromagnetic simulations for single pass beam . Time and frequency domain signals for single pass are presented.

- Analytical analysis for multi-bunch operation

Figure: Analytical simulation for case B, where time and frequency domain simulations are presented. The time domain simulations can be compared to the electromagnetic simulations presented on 1-b. The frequency domain simulations are for multi-bunch beam, repeated at a frequency of 352 MHz.

Figure: Signal power available at the BPM button and BPM electronics as function of the beam current for 5 us pulse, with a 352 MHz bunch repetition rate. The simulation on the left side shows the signal power available in the BPM button and after cables, in dBm.

### BPM signal analysis - Case C

Simulation parameters:

| Case | Chamber diameter (mm) | Button diameter (mm) | Beam energy (MeV) |
| --- | --- | --- | --- |
| C | 100 | 46 | 217 |

  - Electromagnetic numerical simulation

Figure: Electromagnetic simulations for single pass beam . Time and frequency domain signals for single pass are presented.

  - Analytical analysis for multi-bunch operation

Figure: Analytical simulation for case C, where time and frequency domain simulations are presented. The time domain simulations can be compared to the electromagnetic simulations presented on 1-c. The frequency domain simulations are for multi-bunch beam, repeated at a frequency of 352 MHz.

Figure: Signal power available at the BPM button and BPM electronics as function of the beam current for 5 us pulse, with a 352 MHz bunch repetition rate. The simulation on the left side shows the signal power available in the BPM button and after cables, in dBm.

### BPM signal analysis - Case D

Simulation parameters:

| Case | Chamber diameter (mm) | Button diameter (mm) | Beam energy (MeV) |
| --- | --- | --- | --- |
| D | 100 | 46 | 2000 |

- Electromagnetic numerical simulation

Figure: Electromagnetic simulations for single pass beam . Time and frequency domain signals for single pass are presented.

- Analytical analysis for multi-bunch operation

Figure: Analytical simulation for case D, where time and frequency domain simulations are presented. The time domain simulations can be compared to the electromagnetic simulations presented on 1-d. The frequency domain simulations are for multi-bunch beam, repeated at a frequency of 352 MHz.

Figure: Signal power available at the BPM button and BPM electronics as function of the beam current for 5 us pulse, with a 352 MHz bunch repetition rate. The simulation on the left side shows the signal power available in the BPM button and after cables, in dBm.

## BPM electronics performance

The performance of the first prototype of the BPM electronics is presented on this section. Beam phase resolution, position accuracy error and beam position resolution are presented as function of the input power of the electronics. The analysis is presented as function of the RF input power, in units of dBm, instead of as function of beam current due to the many different BPMs and beams conditions at ESS.

Figure: BPM electronics position resolution as function of the RF input power, measured performance. The beam current dependence is also shown since it provides the information about the position accuracy error expected from the BPM electronics as function of the beam current. The calibration factor used for the position calculations is kx = ky = 15 mm.

- References

| [1] | H. Wiedemann, Particle Accelerator Physics, Stanford: Springer, 1999. |
| --- | --- |
| [2] | S. R. Smith, “Beam Position Monitor Engineering,” Argonne National Laboratory, Argonne, Illinois, USA, 1996. |
| [3] | P. Forck, P. Kowina and D. Liakin, “Beam Position Monitors,” Gesellschaft f¨ur Schwerionenforschung GSI, Darmstadt, Germany, 2008. |
| [4] | F. Marcellini, M. Serio, A. Stella and M. Zobov, “Dafne Broad-Band Button Electrode,” Nuclear Instruments and Methods in physics Research, pp. 27-35, 1998. |
| [5] | W. Friesenbichler, Development of the Readout Electronics for the Beam Loss Monitors of the LHC, 2002. |

Electromagnetic simulations performed by DESY.

Electromagnetic simulations performed by DESY.

Electromagnetic simulations performed by DESY.

Electromagnetic simulations performed by DESY.

The RF signal power at the BPM sensor depends on many different parameters such as the BPM sensor type, beam energy, bunch length and beam current. The relation between the beam current and input power were discussed in the previous simulations.