## Slide 1
Infrastructure on Bifrost for Sample Environment
Liam Whitelegg, DTU Physics
Lead Instrument Engineer Bifrost – 27/09/2017

## Slide 2
Bifrost is a cold neutron spectrometer
Positioned in the middle of the West Sector (W4)
Almost all modes will employ the full “white” beam on samples
Aim to study physics of magnetism and superconductivity
5 instrument partners
  DTU (lead)
  KU
  PSI
  LLB
  IFE
Bifrost Instrument Details

## Slide 3
Design of cave driven by;
  Beamline components
  Radiation
  Activation
  Sample environment
  Sample access
Bifrost Sample Area
Rotating Spectrometer Tank
Get Lost Tube
Sample Position
Neutron Guide
Beryllium Collimator

## Slide 4
Very First Design
Stairs up to sample position (3.137m above floor)
Labyrinth entrance with internal crane in cave
Hatch above to crane items in

## Slide 5
Bifrost requires large flux on samples
Experiments could range from 1 to 14 days
Typical experiments could use samples containing Iridium and Lanthanum
Aluminium in sample env. could have a dose rate of 150mSv/hr on contact after 5 minutes of beam.
Could take 30 minutes to access sample position
Necessity for strategy
Activation of Samples & Equip







## Slide 6
TG-2 Design
Lead barrier at sample position
Barrier would have to translate to allow for access
Storage inside cave for active samples and equipment
Internal crane to transport items to sample position
Non-magnetic crane, rad-hard crane

## Slide 7
Estimated equipment usage
  60% VM-1B Magnet
  30% OC 70 (with anvil)
  10% other….
Sample Environment Equipment
VM-1B Magnet
OC 70 – Orange Cryostat
Beam height to flange = 1259mm
Weight ~ 750Kgs

## Slide 8
Weight of dilution stick ~ 50Kgs
Length of dilution stick ~ 1700mm
Hook height required ~ 7.2m
All of this converges on a 9m tall cave
Structural difficulties
Additional racks and supplies
If stored within cave, larger footprint needed
“Dilution Stick Effect”

## Slide 9
Predominant sample access now from roof through top loading
Moved to dual access via platform access
Crane now external to cave
Cave height reduced from 9m to 5.2m
Sample environment auxiliary equipment positioned on roof
New Design
Another picture here

## Slide 10
Lead access hatch linked to PSS system
Ladder access down to platform
Lead lining along base and side to provide shielded access
Secondary access through cave labyrinth to sample position
Personnel Access to Sample
Picture of sample platform

## Slide 11
Several procedures to be developed for varying scenarios
More space to deal with these items
Limits need for internal cave access
Potential for remote, sample removal equipment via crane
Lead lined storage areas and caskets
Guide pins and remote removal possible
Activated Equipment
Pictures & maybe activation table

## Slide 12
Auxiliary Equipment & Services
Sample prep area near kit being used
Danger of radiation damage to components reduced
Space for all auxiliary equipment
No fixed chicanes or patch panels
Aim would be to allow full access to the roof for sample preparation
Service panels all up on the roof
Mobile Filling Dewar
Turbo Pump
Gas Handling System
Magnet Power Supply Rack

## Slide 13
Non-magnetic Requirements
66mT
8.3mT
2.4mT
1mT
Lead and concrete (with SS re-bar) shielding
No magnetic materials within 1m
Specify appropriate motors and position as far from the magnet as possible
All tools to be non-magnetic
Reason | Approx. range (for 15T) Scale as Bm/n | Affected item | Effect
Magnetic forces (unsaturated, dipole approx.) | 0-1m | Cryomagnet, nearby objects | Quench, damage, injury, death
 | 1-2m | Moveable magnetic objects | Pulled towards cryomagnet (see <1m)
 | > 2m | Moveable magnetic objects | Equilibrium positions disturbed
 | <5m | Reed switches (including pacemakers!!!) | Erroneous readings, malfunction
Magnetic forces (permanent magnets/ saturated ferro-magnets) | 0-1m | Motors, encoders, magnetic bearings | Failure, damage