## Slide 1: Beam Monitors Neutron Instrument Division meeting
Fatima Issa
www.europeanspallationsource.se
17 May 2017

## Slide 2: Outline
Market survey of Beam monitors
Integration of Beam monitors into ESS
Parasitic methods for monitoring the beam
Survey on BM requirements from instruments
Facility maintenance and operation

## Slide 3: Beam Monitor
Transmitted beam
Incident beam
scattered
Sample
source
Beam monitor
neutron guide
chopper
Beam monitor
Beam monitor with low attenuation and high transmission factor
Window 1
Window 2
A Little detailed analysis about BM performance is available  Survey what is available on market

## Slide 4: Types of Beam Monitor
Multi-wire proportional chamber (MWPC)
From Mirrotron
From ORDELA
MWPC can be filled with either 3He or 14N gas
From Quantum detector
From LND
Fission Chamber and scintillator
Fission chamber coated with 235U
Scintillator made of LiGlass beads

## Slide 5: Types of Beam Monitor
Multi-wire proportional chamber (2D-MWPC)
Gas Electron Multiplier (2D-GEM)
From Mirrotron
From CDT
0.8 μm boron layer was selected for the sake of the measurement. Thinner layer will be selected for full implementation.
Monitor Manufacturer | 2D-MWPC Mirrotron | 2D-GEM CDT
Active element | 3He | 10B
Total pressure bar | 1.65 | 1
Filled gas | 3He+CF4 | Ar+CO2
Bias voltage (V) | Anode:-3500 Drift: 1500 | -1000
Active area (mm) | 100x50 | 100x100
Window thickness (mm) | 1 | 0.1
Position sensitive BM
G. Croci et al Progress of Theoretical and Experimental Physics, 2014(8):83H01{0, 2014.

## Slide 6: Beam Monitors- Neutron measurement at R2D2
Slit 1
Slit 2
Peak wavelength (Å)
400 2.00
911 0.878
711 1.12
511 1.54
822 0.943
311 2.41
933 0.804
Measuring:
Efficiency
Scattering
Attenuation
Linearity

## Slide 7: Beam monitors-main results
 | MWPC | MWPC | MWPC BM-100X50 | 2D-MWPC | 2D-GEM | Scintillator | Fission chamber
supplier | ORDELA | ORDELA | Mirrotron | Mirrotron | CDT | Quantum detector | LND
Isotope used for neutron capture | 3He | 14N | 3He | 3He | 10B | 6Li | 235U
Gas pressure mbar | Partial pressure 6,0795 | Partial pressure 81,06 | Partial pressure 6,5 | Partial pressure 0.4 | Total pressure 100 | ------ | Total pressure 1013,2
Filled gas | 3He+4He +CF4 | N+CF4 | 3He+CF4 | 3He+CF4 | Ar/CO2 | -------- | P10
Active Area (mm2) | 114 x 51 | 114 x 51 | 100 x 50 | 100 x 50 | Diameter 100 mm | 28 x 42 | Diameter 108.0 mm
Applied voltage (V) | 850 | 850 | 1300 | Anode at -3500V Drift at 1500V | -1000 | 650 | 300
Measured attenuation % | 4.5 | 4.4 | 2.5 | 7.3 | 10.6 | 0.4 | 3.8
Calculated attenuation % | 4.05 | 4.05 | 2.05 | 4.5 | 6.9 | 0.1 | 2.4
Measured Efficiency at 2.4Å | 1.2x10-3 | 2x10-5 | 1.1x10-3 | 1.5x10-4 | 3x10-2 | 5x10-3 | 10-4
Supplier Efficiency at 1.8 Å | 10-3 | 10-5 | 10-3 | 1.5x10-4 | 3x10-2 | 10-4 | 10-4
Scattering (%) | 3.9 ± 0.4 | 3.8±0.4 | 4±0.9 | 9±1 | 10.3±0.7 | 0.74±0.2 | 3.8±0.3
More characterization  monitors will be customized to match ESS requirements

## Slide 8
Horizon 2020 grant agreement 676548
The results of the market survey are summarized here: https://arxiv.org/abs/1702.01037
Submitted to Journal Physical Review Accelerators and Beams (PRAB)

## Slide 9: Outline
Market survey of Beam monitors
Integration of Beam monitors into ESS
Parasitic methods for monitoring the beam
Survey on BM requirements from instruments
Facility maintenance and operation

## Slide 10: Mechanical integration
With this configuration up to 22 Hz no noticeable change in the BM response
The BM should not be in electrical contact with the chopper
Grounding is important see grounding guidelines
BM at the backside of the chopper @Embla

## Slide 11: Beam monitor system Integration
HV integration: Done
BM electronics: In progress
All in EPICS
BM
FE
AMP
DAQ
EF
Timing EPICS
DMSC
ICS Tallinn board
Or DG readout as candidates

## Slide 12: Outline
Market survey of Beam monitors
Integration of Beam monitors into ESS
Parasitic methods for monitoring the beam
Survey on BM requirements from instruments
Facility maintenance and operation

## Slide 13: Parasitic methods for monitoring the beam
Vanadium foil
Timing characteristics are poor
Neutron timing gets blurred out by the distance errors. Not good when accurate peak shapes and timing behind choppers are needed.
Good for looking at the white beam
Info from
Jon Taylor and
Robert bewley (ISIS)
Thanks !!

## Slide 14: Parasitic methods for monitoring the beam
Gamma detector
Incident beam
emitted ϒ-rays
Gamma detector
Using the boron layer on the disc one can know if the beam is open or closed this is done using a gamma detector outside the beam.
A test will be performed this summer being coordinated by the chopper group

## Slide 15: Outline
Market survey of Beam monitors
Integration of Beam monitors into ESS
Parasitic methods for monitoring the beam
Survey on BM requirements from instruments
Facility maintenance and operation

## Slide 16: BM requirements for each instrument
Know each instrument requirements
Reduce the number of options
Select the proper type of BM
Issue guidelines for ESS instruments
Aim: September for draft

## Slide 17: BM requirements for each instrument
Instruments | Answer
LOKI | 
FREIA | 
NMX | 
: document provided
: more information needed
: waiting for answer

## Slide 18: BM requirements for each instrument
Instruments | Answer
TREX | 
VESPA | 
MIRACLES | 
C-SPEC | 
: document provided
: more information needed
: waiting for answer
Other instruments will be sent template soon

## Slide 19: Outline
Market survey of Beam monitors
Integration of Beam monitors into ESS
Parasitic methods for monitoring the beam
Survey on BM requirements from instruments
Facility maintenance and operation

## Slide 20: Facility maintenance and operation Requirements
No access to the bunker just for BM maintenance
Environmental requirement (rate, humidity, temperature,..)
T0 information
Diagnose flux lost
Flux along the beam line
BM system
Based on the market survey and the instrument requirements
Recommendation of BM that full fill each of these tasks
Stability
Cross-Calibration
Radiation hardness
(lifetime >10years @ normal flux)
The model at ESS has very low level of support during operation available : high availability and reliability requirements

## Slide 21
Thank you