## Slide 1: CDR: BEAM CURRENT MONITORS
12-13 July 2017, Bilbao
S. Varnasseri
On behalf of BCM team

## Slide 2
CDR: BCM
OUTLINE:
Location of toroids
Interfaces
Toroid’s functions
ACCT1
ACCT1 installation
Radiation aspects
Combo ACCT2/FCT
Magnetic shielding effectiveness
Combo structure
Installation and integration
Electrical specifications
Verifications and measurements (I)
Verifications and measurements (II)
Delivery by ESS Bilbao
Provisional delivery calendar

## Slide 3
ACCT1
ACCT2/FCT
LOCATION OF TOROIDS

## Slide 4
INTERFACES
CDR: BCM
ACCT1
Beam dump vessel design
Cable length to electronics
EPDM gasket (Vacuum)
External fields from quadrupoles
Electrical specifications
Radiation aspects
Combo ACCT2/FCT
Magnetic Shielding design
External fields from quadrupoles
Flanges tolerances from adjacent devices
Integration within beam line (Vacuum)
Electrical specifications
Cable length to electronics

## Slide 5
ACCT bandwidth (3 dB) | ≥1 MHz
ACCT lower cut-off (3 dB) | 3 Hz
Beam current (nominal) | 80 mA
Electronics full scale range | ±10 V
FCT bandwidth | >352 MHz
FCT readout rise/fall time | <10 ns
Main measurement requirements of toroids
TOROIDS FUNCTIONS:
  ACCT1, ACCT2: MEBT beam current, macropulse
  FCT: Chopper rise/fall time, Beam fast timing characteristics
CDR: BCM

## Slide 6
Quadrupoles field data |  | 
Q4 Gradient | 5 T/m | +20%
Q5 Gradient | 6.5 T/m | +20%
ACCT1 distance from Q4 | 373 mm | 
ACCT1 distance from Q5 | 339 mm | 
External magnetic field sources for ACCT1
The external magnetic fields generated by quadrupoles are far lower than ACCT1 toroid limit.
Radial: 2 mT max
Transverse : 10 mT max
ACCT1
CDR: BCM

## Slide 7
ACCT1 Location
CDR: BCM
Calibration
Signal

## Slide 8
RADIATION ASPECTS
Since the ACCT1 is located just after beam dump, the radiation aspects of the toroid has been investigated. The expected radiation dose calculation in the document ESS-0087737 )
ACCT Component | Material | Radiation Resistance (Gy) | 
Wiring Insulation | PVC | 2x105 | Ok
 | Fiber Glass | >108 | Ok
 | +Rubber Adhesive | >106 | Ok
Stress Absorbent | Silicon Rubber Tape SIR | 5x105 | Ok
 | Silicon Rubber SIR | 2x105 | Ok
BNO | ETFE Tefzel | 106 | Ok
SMA (Rad-Hard) | PEEK | 6x107 | Ok
SMA | PTFE | 103 | Not used
CDR: BCM

## Slide 9
Combo ACCT/FCT
External field data |  |  | 
Q10 (negative or Vertical) | -31.5 T/m | 220 A (nominal) | 300 A
Q11 (positive or Horizontal) | 24.02 T/m | 168 A (nominal) | 300 A
Distance Combo-Q10 (center to center) | 100 mm |  | 
Distance Combo-Q11 (center to center) | 100mm |  | 

## Slide 10
S-Iron
Mu-Metal
Toroids
Toroids
MAGNETIC SHIELDING EFFECTIVENESS
Magnetic field (T)
CDR: BCM
Magnetic field outside and after attenuation in the toroids region

## Slide 11
COMBO STRUCTURE
CDR: BCM
Pre-CDR comment

## Slide 12
SOLUTION TO THE COMMENT RID#1:
The gap has been removed. And a reinforced welding is done to to ensure that the flange weight can be handled once the assembly will be mounted in the beam line.
CDR: BCM

## Slide 13
CDR: BCM

## Slide 14
Sensitivity | 2.5 V/A
Turns ratio | 10:1
Rise time | 585 ps
Droop | < 10 %/us
Upper cutoff frequency (-3dB) | 600 MHz
Max. charge/pulse (pulses <1ns) | 0.2
Full scale range | ±80 mA (±10V)
Lower cutoff frequency (-3dB) | 3Hz
Droop | < 2%/ms
Upper cutoff frequency (-3dB) | 1 MHz
Rise time | 350 ns
Combo (FCT)
ACCT1/ACCT2
ELECTRICAL SPECIFICATIONS
CDR: BCM

## Slide 15
INSTALLATION AND INTEGRATION: BCT LOCATIONS ON RAFTS
CDR: BCM

## Slide 16
RAFT09 AND RAFT04 INTEGRATION

## Slide 17
RAFT09 INTEGRATION 1/4
Assembly of Shielding
CDR: BCM
COMBO TRANSFORMER

## Slide 18
RAFT09 INTEGRATION 2/4
Integration of Scrapper chamber and vacuum system.
Fulfill vacuum and position requirements.
CDR: BCM

## Slide 19
RAFT09 INTEGRATION 3/4
Disassembly of vaccum system and integration of quadrupoles.
CDR: BCM

## Slide 20
RAFT09 INTEGRATION 4/4
Final installation of vaccum system and quadrupoles.
CDR: BCM

## Slide 21
RAFT04 INTEGRATION 1/2
Assembly of Beam Dump.

## Slide 22
RAFT04 INTEGRATION 2/2
Same integration steps of Raft09.
CDR: BCM

## Slide 23
VERIFICATION AND MEASUREMENTS(1)
Verifications for ACCT1
Frequency response from 5 Hz to 3 MHz.
Time response to a step to measure the ACCT output rise time.
Time response to a long pulse, typ. 30 us.
Differentiation measurement to measure the droop of the ACCT
Noise measurement
Verification for ACCT2
Frequency response from 5 Hz to 3 MHz.
Time response to a step to measure the ACCT output rise time.
Time response to a long pulse, typical 30 us.
Differentiation measurement to measure the droop of the ACCT
Noise measurement
CDR: BCM

## Slide 24
Verifications for FCT
Frequency response over the full bandwidth
Time response to a short pulse, typical 250 ps rise time.
Time response to a step
Differentiation measurement to measure the droop of the FCT
VERIFICATION AND MEASUREMENTS(2)
CDR: BCM

## Slide 25
Item | Description | Comments | References in technical design document MEBT-BI-BC91-04
ACCT1 | In-Air, ID82 | Installed within Beam Dump | Section 5.1 Section 3
ACCT2/FCT | Combo In-Flange | Includes the mu-metal layers, soft iron layer and two toroids ACCT and FCT with CF flanges | Section 5.2 Section 3
ACCT1-E cable | 20m, Mulrad-2 | BNO connectors | Section 6.2.1
ACCT2-E cable | 20m, Mulrad-2 | BNO connectors | Section 6.2.1
ACCT1 Analogue electronics | Full scale: ±80mA Output: ±10 V | In: BNO female Out: BNC female | Section 3
ACCT2 Analogue electronics | Full scale: ±80mA Output: ±10 V | In: BNO female Out: BNC female | Section 3
DELIVERABLE BY ESS BILBAO
CDR: BCM

## Slide 26
Item | Start | Finish
Finish design (including Mechanical) | 15- July- 2016 | 15- April- 2017
Detailed drawings (Joint Bilbao with Bergoz) | 02- May- 2017 | 15- June- 2017
Review of the Technical Report | 1- June-2017 | 15-June-2017
Technical Tender Document | 20- July- 2017 | 30- August- 2017
Tender offer/Order | 17- Oct- 2017 | 
FAT | 21-Feb- 2017 | 
Delivery to ESS Bilbao | 12-March- 2017 | 
Verification at ESS Bilbao | 2-April- 2018 | 19- April- 2018
Documentation | 2-May- 2018 | 30- May- 2018
Ready for Installation | June- 2018 | 
PROVISIONAL DELIVERY CALENDAR
CDR: BCM