## Slide 1: Status of the ESS MEBT EMU
EMU Team
ESS-Bilbao
Beam Instrumentation Group
July – 2017
Notes: Decir que hablo en parte del equipo de ESS-Bilbao
Decir que es básicamente el trabajo presentado para el CDR.

## Slide 2: ESS MEBT
SLIT
GRID

## Slide 3: L4 Requirements
User Defined ID | DOORS-ID | Name | Description
 |  | MEBT transverse phase space measurement | The beam distribution in transverse phase space shall be measured in the MEBT section. The measurement can be invasive.
MEBT-L4-PBI-110 | MEBT.PBI-54 | MEBT transverse phase space measurement planes | The transverse phase space measurement shall consist of one horizontal and one vertical measurement. Both measurements shall independently fulfil all requirements.
MEBT-L4-PBI-120 | MEBT.PBI-55 | MEBT transverse phase space 95% emittance measurement error | The transverse emittance containing 95% of the beam shall be measured with a total measurement error of less than ±10 % of the measured value.
 |  | MEBT transverse phase space RMS emittance measurement error | The transverse RMS emittance shall be measured with a total measurement error of less than ±10 % of the measured value.
 |  | MEBT transverse phase space measurement dynamic range | The transverse phase space measurement shall have a dynamic range of 1000.
Notes: Invasive measurement of phase space
Independently measure horizontal and vertical BUT NOT at the same time!
The 95% and RMS emittance reconstruction within ±10%
Dynamic range 1000

## Slide 4: EMU Measurement

## Slide 5: EMU Measurement

## Slide 6: Main parameters
These are the main parameters that affect to the reconstruction of the emittance
Distance Slit – Grid: 362.9 mm
Slit aperture: 100 µm
Slit thickness: 200 µm
Grid wires: 24
Wires diameter: 35 µma
Grid pitch: 500 µm
Expected maximum signal ~100 µA
Notes: Parameters coming from Benjamin’s design report

## Slide 7: Error assignment
The error in the emittance reconstruction depends in the Slit – Grid distance. For 400 mm the main contribution to the error are:
  Error in the position of slit and grid ~2%
  Error due to scattering (200 µm thickness slit) ~1%
  Error due to space charge ~1.5 %
Sum of errors ~5%
We still have some margin for the 10% goal
Notes: Data comming from Benjamin’s report

## Slide 8: Slit design
Error | Description | Responsable
∆αfg | Angular error between the fixed flange of the actuator (f) and the linear guides (g) | Contractor
∆αgm | Angular error between the guides (g) and the mobile flange of the actuator (m) | Contractor
∆αme | Angular error between the mobile flange (m) and the shaft (e) | Contractor
∆αes | Perpendicularity between the shaft (e) and the Slit apperture (s) | Contractor
∆αfs | Total angular error between the fixed flange of the actuator and the reference plane of the Slit aperture | Contractor
Notes: Allowed mechanical errors
0.5 degree of angular precision in all directions (slit aperture and flange)
Slit aperture 100 microm +- 20 microm

## Slide 9: Mechanical requirements
0.5 degree of angular precision in all directions (slit aperture and flange)
Slit aperture 100 µm ± 20 µm

## Slide 10: Slit subsystem
Actuator
Slit arm
155 mm stroke
WS design as baseline

## Slide 11: Slit arm
A – Graphite plates; B – Substrate; C – Cooling tubs D – Shaft; E – Shaft holding piece
A
B
Slit head
D
E
C

## Slide 12: Slit Shaft
Welled to the top holding piece
Screwed in the bottom part

## Slide 13: Slit Actuator
Same design as WS
A
C
D
I
K
J
B
E
F
G
H
O
R
Q
M
N
L
P
S

## Slide 14: Slit design
B
A
Two Graphite plates – A chamfer to get 200 µm thickness
50 µm rabbet to get 100 µm aperture
Chamfer

## Slide 15: Graphite Plates

## Slide 16: Slit Thermomechanical design
FEM model

## Slide 17: Slit Thermomechanical design
Deformation

## Slide 18: Slit Thermomechanical design
Results for Stainless Steel (250 N contact force)
Case | ΔT (ºC) | σ1 (MPa) | σ3 (MPa) | σInt (MPa) | σInt/σLim. | Slit. Def. (μm)
Initial | 17 | 0 | 0 | 0 | 0 | 0
Steady | 57 | 0.75 | -1.5 | 1.5 | 2 % | 2
Transient | 463 | 7.6 | -26 | 27 | 32% | 2.3
Total (Init + Steady + Transient) | 537 | 7.6 | -26 | 27 | 32% | 4.3

## Slide 19: Assembly
A
B
C
D
Notes: A – Weld channel
B – Assembly actuator components: support, below, flange
C – Weld new section of cooling tubes
D – Up vacum feedthrough, inssert shaft, screw from interior and lower feedtrhugh Then, Weld cooling tuve.

## Slide 20: Slit Cooling
The refrigeration circuit flexible tubing:
  6 mm fitting in one side
  adaptor to a 6 mm pipe on the other side,
The cooling feedthroughs:
  6 mm tube fitting in one side
  Female tapered cylindrical BSP connector in the other side (Cooper gasket)

## Slide 21: Slit Cabling Path
C
A
B
D
C
D
A
B
E
Cables to the Patch Panel (C):
Limit switches from switches position (D)
Motor and brakes from motor position (A)
Encoder cable from encoder position (B)
Patch Panel:
Motor and brake
travel limit switches
MPS Limit Switch (LEMO)
Encoder
Connector to fix cables and protection

## Slide 22: Grid Subsystem
Actuator
X
Y
Z
Grid arm
80 mm stroke
CF40

## Slide 23: Grid arm
E
A
C
Grid Head
B
D
A – Wire PCB; B – Bias plates; C – Alignment system D – Shaft; E – Shaft holding piece

## Slide 24: Grid head
A – Wire PCB; B – Bias plates; C – Screws for bias wires D – Signal connector; E – Alignment system,
F – Fork, G – Screws, H – Holes for precision pins
A
B
C
D
E
F
H
G

## Slide 25: Alignment system
A
B
C
A – Precision Pin; B – Screw ; C – Guide
(system similar to WS)

## Slide 26: Wires PCB design
24 Tungsten wires
35 µm diameter
~50 mm long
500 µm pitch
Covers 11.5 mm
Glued with H20E Epoxy
24 pin vacuum connector

## Slide 27: Grid thermomechanical design
Parameter |  | Value | 
Slit-Grid Distance |  | 400 | mm
Slit | σx | 3.30 | mm
 | σx’ | 6.65 | mrad
 | σy | 3.98 | mm
 | σy’ | 10.08 | mrad
 | I0 | 62.5 | mA
Grid (Slit H.) | σx | 6.37 | mm
 | σx’ | 7.12 | mrad
 | σy | 0.35 | mm
 | σy’ | 0.89 | mrad
 |  | 0.56 | mA
Grid (Slit V.) | σx | 0.46 | mm
 | σx’ | 1.165 | mrad
 | σy | 0.354 | mm
 | σy’ | 10.76 | mrad
 |  | 0.68 | mA
Parameter |  | Value | 
Slit-Grid Distance |  | 400 | mm
Slit | σx | 3.30 | mm
 | σx’ | 6.65 | mrad
 | σy | 3.98 | mm
 | σy’ | 10.08 | mrad
 | I0 | 62.5 | mA
Grid (Slit H.) | σx | 6.37 | mm
 | σx’ | 7.12 | mrad
 | σy | 0.35 | mm
 | σy’ | 0.89 | mrad
 |  | 0.56 | mA
Grid (Slit V.) | σx | 0.46 | mm
 | σx’ | 1.165 | mrad
 | σy | 0.354 | mm
 | σy’ | 10.76 | mrad
 |  | 0.68 | mA

## Slide 28: Grid thermomechanical design
Temperature evolution
Temperature distribution
Temperature below thermionic limit < 1500 K
Wire Pretension > 40 g (0.4 N)

## Slide 29: Wires PCB Prototype
Distance between wires: 500 ± 50 µm
100 g of preload tension

## Slide 30: Grid actuator

## Slide 31: Small differences with Slit actuator
A

## Slide 32: Cabling
Bias wire goes through the screws
Signal cables go through the connector

## Slide 33: Grid cabling path
C
D
A
B
E
A
C
E
B
F
D
Cables to the Front-End:
Signal cable
Cables to the Patch Panel (C):
Limit switches from switches position (D)
Motor and brakes from motor position (A)
Encoder cable from encoder position (B)
Patch Panel:
Motor and brake
travel limit switches
MPS Limit Switch (LEMO)
Encoder
Connector to fix cables and protection

## Slide 34: Acceptance
Vacuum test
  Test each subsystem to the required vacuum + rga
Metrology test
  Ensure fundamental measurement are OK
Hydraulic test
  Test the slit cooling circuit for leakage
Actuator verification

## Slide 35: Verification
Motors & control
Analog FE & control
Complete system
  Cross-talk
  Integrity check

## Slide 36: Schedule
Milestone | Delivery Term
Contract Formalization | T0
Delivery of Blueprints | T1 = T0 + 1 month
Acceptance of Blueprints | T2 = T1 + 1 month
Manufacture and delivery of EMU-H and Verification Documents | T3 = T2 + 3 months
Acceptance of EMU-H | T4 = T3 + 1 month
Manufacture and delivery of EMU-V and Verification Doc. | T5 = T4 + 5 months
Acceptance of EMU-V | T6 = T5 + 1 month
Vertical Test | T7 = T6 + 3 months
Total Duration of the Project | 15 Months

## Slide 37: Risks
Id | Event | Cause | Impact | Treatment plan | Responsible for risk treatment | Status
MEBT-PBI-EMU-01 | Problems with postioning the wires. Incorrect distance between wires | Dedicated experience is mandatory | Precision in the measurements | 1- Find partners 2- Build a prototype in advance | BIO | Treated
MEBT-PBI-EMU-02 | Wire breaks | Handling during assambly/instalation | Schedule delay, failure at the integration | Consider spares in budget | BIO | Teated
MEBT-PBI-EMU-03 | Wire breaks | During operation in commissioning | Schedule delay, failure at the integration | Incorporate “wire integrity” check | BIO | Teated
MEBT-PBI-EMU-04 | Grid Actuator specification not meet | Misunderstanig with contractor about specification | Precision in the measurements | Elaborate MEBT-BI-EM90-06 with specifications, and follow up. | BIO | Teated
MEBT-PBI-EMU-05 | Grid Actuator Missalignemnt (tilt) | Impossibility to meet tolerances by contractor | Precision in the measurements | Design Grid with alignment screw | BIO | Teated
MEBT-PBI-EMU-06 | FE not meeting design goals | Inadequate design | Precision in the measurements | Build and test prototype | BIO | Treated
MEBT-PBI-EMU-07 | FE radiation sensitive |  | Integrity of device is under risk | Request ESS a dedicated bunker in the tunnel | ESS | Treated

## Slide 38: Risks
Id | Event | Cause | Impact | Treatment plan | Responsible | Status
MEBT-PBI-EMU-08 | Noise in the signal from secondary electrons | Not handling SE particles | Precision in the measurements | Integrate a bias voltage +-1kV in the design | BIO | Treated
MEBT-PBI-EMU-09 | Slit graphite not meeting specification | Misunderstanig with contractor about specification | Precision in the measurements | Follow manufacturing with bi-wekly report | BIO | Started
MEBT-PBI-EMU-10 | Cross-talk between signal wires | Vacuum cables have no enough insulation | Precision in the measurements | Measure cross-talk in sample cable/Study alterative design | BIO | Started
MEBT-PBI-EMU-11 | Graphite plates break | Handling during assambly/instalation | Schedule delay, failure at the integration | Allocate budget for spare parts | BIO | Started
MEBT-PBI-EMU-12 | Slit Actuator specification not meet | Misunderstanig with contractor about specification | precision in the measurements | Elaborate MEBT-BI-EM90-06 with specifications, and follow up. | BIO | Teated
MEBT-PBI-EMU-13 | Development in VME | Restricted market for µTCA | Schedule delay, failure at the integration | Develop plan to migrate to µTCA Confirm delivery dates for HW and SW | ESS | Started
MEBT-PBI-EMU-14 | Migrate to new framework for GUI | New framework not finished yet | Schedule delay, failure at the integration | Clarify migration responsibility | ESS | Started