## Slide 1: Multi-Grid: the New Generation Large-Area Detector for Neutron Scattering Science
Anton Khaplanov
On behalf of ESS Detector Group and Collaborators
www.europeanspallationsource.se
July 13, 2017

## Slide 2: Outline
B10 detector development
Multi-Grid detector
  Where it will be used
  How it works
  Characterization tests
  Current development
Multi-Grid test at CNCS, SNS
  Setup at CNCS
  Results from CNCS
Conclusion

## Slide 3: Acknowledgements and Publications
ILL:
Bruno Guerard, Jean-Claude Buffet,
Jean-Francois Clergeau, Anthony Leandri ,
Victor Buridon, Fabien Lafont
ESS:
Anton Khaplanov, Fatima Issa, Richard Hall-Wilton, Oliver Kirstein, Tomasz Brys, Michail Anastasopoulos, Isaak Lopez Higuera, Richard Bebb, Sara Arranz, Carina Höglund*, Linda Robinson*, Susan Schmidt*
Centre for Energy Research (Hungary):
Eszter Dian
Linköping University:
Jens Birch, Lars Hultman, (also *)
SNS:
Ken Herwig, Georg Ehlers, Michelle Everett, Kevin Berry
Earlier – the participants of the CRISP project on Large-Area detectors.
Earlier publications:
B4C layers:
*C. Höglund et al, J of Appl. Phys. 111, 104908 (2012)
Characterization:
*A. Khaplanov et al., arXiv:1209.0566 (2012)
*B Guerard et al., NIMA, 720, 116-121 (2013), http://dx.doi.org/10.1016/j.nima.2012.12.021iJ
*J. Correa et al., Trans. Nucl. Sc. (2013), DOI: 10.1109/TNS.2012.2227798
*A. Khaplanov et al., (2014) J. Phys.: Conf. Ser. 528 012040 doi:10.1088/1742-6596/528/1/012040
Gamma sensitivity:
*A. Khaplanov et al., JINST 8, P10025 (2013), arXiv:1306.6247
Alpha background:
*A. Khaplanov et al., JINST 10, P10019 (2015); doi:10.1088/1748-0221/10/10/P10019
Latest publication:
A.Khaplanov et al. “Multi-Grid Detector for Neutron Spectroscopy: Results Obtained on Time-of-Flight Spectrometer CNCS" https://arxiv.org/abs/1703.03626 2017 JINST 12 P04030
WP 4.3: Large-Area Detectors
Horizon 2020 grant agreement 676548

## Slide 4: B10 Detectors as Alternative to He3
3He crisis, demand vs. supply:
Challenge, but also a new opportunity:
In addition to lower cost,
Improve rate performance
Improve position resolution
No requirement for gas purification
No high pressure gas
Several other isotopes have capture reactions that allow thermal neutron detection. We use:
10B + n  7Li + α + 2.97MeV
NATB contains ~20% 10B – a relatively inexpensive isotope.
Still working with gas detectors, but now the converter is solid
B-10 film
substrate
n
conversion
gas
7Li or α
7Li or α
pressure vessel
n
conversion
3He gas
3H or p
3H or p
3He detector
10B detector

## Slide 5: B10 Detectors as Alternative to He3
B-10 film
on both sides
Al substrate
n
conversion
gas
fragment 1
fragment 2
gas
Anode wires
n
fragment 1
fragment 2
n
fragment 1
fragment 2
Anode wires
B-10 film
on 1 side
Multi-Grid:
Multiple layers at normal incidence
Multi-Blade:
Single layers at ~5° incidence
140mm
100mm
n
Multi-Blade
reflectometry
demonstrator
Many other geometries are possible
*F.Piscitelli, JINST 12, P03013 (2017)
See presentations J. Lacy ThB1-2 and G. Nowak ThB1-3
~5% efficiency
in 1 layer

## Slide 6: Multi-Grid: How it Works
n
Low cost for large area
Modular
High total efficiency
Many layers – lower local rate
No need for resistive readout – low gain
n
n
Introduced at ILL, jointly developed by ILL and ESS under CRISP project,
now under BrightnESS

## Slide 7: Application of Multi-Grid: Time-of-Flight Spectrometers
Time-of-flight information
2-5 m sample to detector distance
 Energy transfer
20-40 m2 detector area
Position resolution O(1inch)
Momentum transfer
High dynamic range
Low background
n Ei, Qi
n Ef, Qf
Incident beam
Sample
Detected ToF spectrum
Detected image
IN5 at ILL

## Slide 8: Time-of-Flight Spectrometers at ESS
Instrument | C-SPEC | T-REX | VOR
Sample – detector | 3.5 m | 3 m | 3 m *
Detector height | 3.5 m | 2.2 m | 3 m *
Detector area | 29 m2 | 21 m2 | 37 m2 *
First 15 instruments chosen
3 ToF spectrometers for ESS
CSPEC and T-REX – design phase
VOR – accepted proposal
Multi-Grid as baseline detector technology
*as proposed

## Slide 9: CSPEC at ESS
Cold spectrometer
0.2 meV < Ei < 20 meV
29m2 detector
Horizontal coverage 5° to 135°
Vertical coverage -25° to 25°

## Slide 10: T-REX at ESS
Thermal / bispectral spectrometer
2 meV < Ei < 160 meV
21m2 detector
Horizontal coverage -36° to -1° and 1° to 144°
Vertical coverage -25° to 15°

## Slide 11: Multi-Grid Prototypes for Characterization
Measurements:
Absolute efficiency
Uniformity
Position reconstruction
γ sensitivity
Rate tests
Ageing
MG.24 @R2D2 @IFE
Proto 2 at CT2 @ILL
Proto 1 at CT1 @ILL

## Slide 12: Efficiency measurement at Test Beam
Proto 2 at CT2 @ILL
Efficiency vs. layer thickness
n beam
Counts compared to a carefully calibrated reference detector,
in a constant beam,
background subtracted
 error bars below 1%
ILL “Honeycomb” detector

## Slide 13: Source Measurements
A
B
C
D
4 regions of grids assembled in different ways
(edge effects due to incomplete scanning range)
Source motion
Laboratory sources:
Thermal neutrons: α-Be or 252Cf +PE moderator;
γ sources: 57Co, 60Co, 133Ba, 137Cs, etc.
Source Testing Facility, Lund
Scanning table, ILL
Fast neutron sources: α-Be or 252Cf;
Water tank with beam ports

## Slide 14: First Multi-Grid tested on an Instrument
Detector tested on cold chopper spectrometer IN6 @ILL
Active Area 30 x 45 cm2
96 grids in 6 modules of 16 grids, 360 anode wires
Operated for ~2 weeks, replacing 25 He3 tubes
@ 4.1, 4.6 and 5.1 Å incoming wavelength
Main topics of the test:
First ToF spectra
Background investigation
Elastic line shape
*A. Khaplanov et al. 2015 JINST 10 P10019
*A. Khaplanov et al. J. Phys: Conf Ser 528 (2014) 012040

## Slide 15: Intrinsic background
Spectrometers work with a very large dynamic range  very high signal-to-noise is required.
Intrinsic background often found in Al due to α emitter impurities (U and Th).
Spectrum from standard Al
N + 10B7Li + α
7Li α
α decay U, Th
x50-x100 background reduction using:
high-purity Al
Ni-plating
Ni-plating on a 0.5mm Al blade
A. Khaplanov et al. 2015 JINST 10 P10019

## Slide 16: γ Sensitivity Investigated
* A. Khaplanov et al., JINST 8, P10025 (2013), arXiv:1306.6247
Operation HV, can choose γ sensitivity 10-6 – 10-9
High statistics 137Cs measurement
Plateau for γ and n measurement
Beam + source setup for γ sensitivity measurement
γ rejection demonstrated in ToF @IN6

## Slide 17: IN5 Demonstrator – Large-Scale
Detector:
2.4 m2 active area
Blade x18432:
enriched B4C coating
good adhesion, uniformity,
Module x8:
wire-frame coincident readout
Grid x1024:
low activity,
minimal dead material
80cm
3 m

## Slide 18: 10B4C coating facility in Linköping
ESS workshop, Linköping, Sweden
10B4C coatings for detectors
Using DC magnetron sputtering
B4C layer analysis
at Linköping University:
Thickness, uniformity
Composition, impurities:
Produced blades for MG
Many other types of substrates can be coated

## Slide 19: New grid design at ILL
New Grid design at ILL
3x grid width – fewer dead gaps
24 layers - higher efficiency
Coating on parallel blades – absorption of scattered neutrons
Curved grids
n

## Slide 20: Gas pressure study at ILL
Development of gas circulation system
Detector can be operated at
50 mbar – 1 bar
Ambient vacuum
Performance as a function of pressure – the detector can be operated at 50 mbar with no loss of performance.

## Slide 21: Test on CNCS, Cold Neutron Chopper Spectrometer
A kind offer by K. Herwig to test MG at SNS
Recommendation of 2015 ESS annual review
Goals:
Test at spectrometer
Side-by-side comparison to He3
In parallel with user experiments
Dedicated tests
Solution:
Size = half of “8-pack” module – 1.1m x 19cm
Installation June-July 2016
Operated July 2016 to June 2017

## Slide 22: Design and Construction of MG.CNCS
1.1m

## Slide 23: Multi-Grid installed at CNCS
Detector installed at 57° scattering angle
Inaccessible between summer and winter shutdowns
DAQ setup outside, accessible remotely
MG
He-tubes
3He 8-packs
M-G
1.1m
19cm
20cm
View from sample
2m

## Slide 24: Methods Used for Tests
Vanadium
Counts over time, background
Single crystal
High local rate
Fast neutrons (fission, prompt)
Water, Ice
Allowing to measure:
Energy resolution – efficiency, scattering – background sensitivity – saturation
Gamma background
(NaI detector)
x2 lower fast n counts in MG
Rate and background analysis

## Slide 25: Elastic Energy Resolution

## Slide 26: Efficiency
Counts from integral of the elastic peak in dE
Efficiency measured relative to the nearest He3 tubes
Error bars: 5-8%, agrees with prediction

## Slide 27: Single Crystal Reflection
Crystal
Incident beam
Reflected beam at Ei = 17.20meV
Reflected beam at Ei = 13.74meV
A single crystal reflects neutrons according to Bragg’s law:
Resulting in an intense spot seen by detector
He3
MG

## Slide 28: Single Crystal Reflection
He3
Multi-Grid
Image recorded in each detector as a function of time during elastic peak
Each frame advances 30 μs
No loss of position resolution or saturation observed in Multi-Grid

## Slide 29
GEANT4 Simulation of MG.CNCS
Etrf = Einitial - Efinal
Elastic peak
Scattered background from sample environment
Scattered neutron background from detector
Detector
Sample environment
3.3 m
Isotropic scattering particles simulated from sample
Simulation of MG.CNCS
Neutrons generated at sample
Reconstructed energy calculated from ToF
Sample environment included
Detector tank window included
Contribution of each component can be separated
*See presentation by K. Kanaki , MoF1-4

## Slide 30: Conclusion
Multi-Grid developed and characterized.
Large-Area detectors possible again
Operated at CNCS over 11 months.
Multi-Grid baseline detector for
    CSPEC and T-REX at ESS
Instrument-focused design is underway
Proposal for a thermal instrument test in preparation
Thanks to the colleagues at SNS for this opportunity and cooperation!
WP 4.3: Large-Area Detectors
Horizon 2020 grant agreement 676548
2017 JINST 12 P04030
(https://arxiv.org/abs/1703.03626 )

## Slide 31: Extra slides

## Slide 32: Time-of-Flight Spectrum
Elastic peak
Neutron energy loss
Neutron energy gain
Prompt peak
Time between source pulses

## Slide 33: Energy Reconstruction
Energy transfer: Ef-Ei
Raw ToF spectrum:
no information on geometry used
dE spectrum: Including information on position of each detected neutron.

## Slide 34: 10B4C coatings
Coatings of 10B-enriched B4C thin films have been produced for a variety of detectors and substrate materials
Some new challenges:
Temperature sensitive substrates
No substrate bending when 1-side coated
Pre-defined coating resistivity
Complex substrate shapes
Publications:
*C. Höglund et al, J of Appl. Phys. 111, 104908 (2012)
*S. Schmidt et al, J. of Materials Science 51, Issue 23 (2016)
*C. Höglund, Rad. Phys. and Chem. 113 , 14-19 (2015);

## Slide 35: Single Crystal Reflection
He3
MG
The 6th image frame from previous slide
(Peak rate in 1 tube / 1 line of pixels ≈ 150kHz)
Position:
Charge division in He3 PSD – saturation causes loss of position resolution
Coincidence between grids and wires in Multi-Grid
Local count rate limit:
space charge in gas (*)
  x10 lower in MG per event
  x10 lower in MG due to lower gain
Dead time of electronics
depends on electronics, similar per channel
No saturation observed in Multi-Grid

## Slide 36: Aluminium Bragg Edge
Incident n
Sample
Sample environment
Detector
Scattering effects from SE and detector

## Slide 37: Efficiency Optimizations
Optimized grid
1.5μm
0.5μm
1.0μm
n
Layer thicknesses in MG.CNCS (16 blades):
7 blades 0.5μm,
7 blades 1.0μm,
3 blades 1.5μm
Proposed for a thermal detector (20 blades):
4 blades 1.0μm,
10 blades 1.25μm,
6 blades 2.0μm.
Efficiency as a function of wavelength for MG optimized for cold vs. thermal spectrum
Cold – optimization centered on 4Å
Thermal – optimization centered on 1Å
Thermal optimization
Thermal Optimization
Tested at MG.CNCS
*F. Piscitelli, P. van Esch, 2013 JINST 8, P04020