## Slide 1: BEAM INSTRUMENTATION
Tom Shea
and the ESS Beam Instrumentation Team

## Slide 2: Outline
Accelerator and its Beam
Beam Instrumentation by Function
  Measurements
  Protection
Summary

## Slide 3: Accelerator Layout
Spokes
0.4 µs
Medium Beta
0.4 µs
High Beta
0.7 µs
MEBT
beam
DTL
0.5 µs
The Beam
Long pulse proton beam 2,86 ms
Pulse frequency 14 Hz
Proton energy 2 GeV
Beam power, time averaged 5 MW, within pulse 125 MW
Beam current 62,5 mA
Beam Instrumentation
Sensory system
Distributed throughout the accelerator
Measures beam beam properties
Delivers measurements to the control system
Ion Source
Target
75 keV
3.6 MeV
90 MeV
216 MeV
571 MeV
2 GeV

## Slide 4: Beam: Phase Space, Emittance
Transverse: x (mm), x’ (mrad)
Longitudinal: z, z’
(also: eV, seconds)
Transverse: y (mm), y’ (mrad)
Transverse: x (mm), y (mm)
Eshraqi

## Slide 5: ESS Beam Parameters at Target
Beam footprint enclosing 97.5% beam fraction | 180 mm (horizontal) 60 mm (vertical)
Nominal time-averaged peak current density | 56 µA/cm^2
Max displacement of footprint from nominal position | ±5 mm (horizontal), ±3 mm (vertical)
Nominal RMS size of Gaussian beamlet | 16 mm (horizontal) 5.3 mm (vertical)
Raster frequencies | 40 kHz (horizontal) 29 kHz (vertical)
Sector of Target Wheel
Gaussian beamlet

## Slide 6: Beam Instrumentation Devices
Beamline device
Signal Processing
Buffering
Shielding
configuration
beam interlock
Machine Protection System Interfaces:
post mortem
cabling
device state
device health
Accelerator Interfaces:
Controls network
Timing/event
RF reference
Low-latency data
Possible Functions:
Measurements
  Beam Accounting
  Measurement of Beam Centroid
  Measurement of Beam Distribution
Protection
Notes: Only those most relevant for MPS
Also include optional out of band config for thresholds and direct MPS link to beamline device state (like SNS)
also, environmental considerations: electromagnetic interference, radiation, cryogenic environment, ultra-high vacuum, etc.
Reference: Forck lecture notes

## Slide 7: Beam Accounting
Measuring beam current and where it is lost
  Constituents of beam extracted from ion source (what fraction is protons?)
  Lost protons that damage or activate accelerator
  Beam leaving aperture and overheating target/dump components
Systems
  Doppler
  Faraday Cups
  Beam Current Monitors
  Beam Loss Monitors
    Neutron detectors
    Ion chambers
  Aperture Monitors
Beam pulse parameters for diagnostics measurements extend from a few mA for a few µs at low repetition rate up to full duty factor.

## Slide 8: Beam Accounting: Doppler and Faraday Cup
Doppler system: optical spectrometer measures doppler shifted visible light emission and reports ion species
Faraday cup measuring 75 keV beam
as installed @Catania
Several milliseconds
Current vs. time
Beam into conducting cup

## Slide 9: Beam Accounting: Beam Current Monitor
Single channel measurements:
1% accuracy
<1 µs response to measure current and remove beam permit on
  Wrong pulse parameters
  Difference from previous pulse
Difference measurements between two devices:
Digital optical link and processing
Minimum threshold: ~2% current loss.
< few µs typical to calculate current difference and remove beam permit on fast loss
Forck, Grishin, Hassanzadegan
Current vs. time

## Slide 10: Beam Accounting: Beam Loss
det2 “detector” volume
x
y
z
det1
det0
I. Kittelmann
Absorbed dose map, all particles, from point loss @ 2 GeV
Neutron hit maps for 3 different localized loss locations along the DTL tank1: possible correlation of the distribution peak position with the loss location
M. Jarosz

## Slide 11: Beam Accounting: neutron Beam Loss Monitors
SLOW
FAST
Recoil protons produced by neutrons in polypropylene
High flux high energy n's (>0.1 MeV)
Faster response
Detector size
~30x30x30cm3
~3kg
Easy to transport
Gas chamber
Different thickness of polyethylene can be added
(n, α) 10B reaction
Detection of fast neutrons after moderation in polyethylene (~4cm)
More efficient, 4π, but slower response
CEA Saclay

## Slide 12: Beam Accounting: Ion Chamber Beam Loss Monitors
Particle shower from lost beam ionizes gas in the Ion Chamber
Electric field sweeps charge to electrodes
Resulting current is measured
Current sensitive digitizer
Shortest pulse (5 us):
 1% of 1 W/m beam loss
Pico Amps

## Slide 13: Beam Accounting: Aperture Monitors
Aperture Monitor concept:
Nickel blades and K-type thermocouples
Instrumentation Plug
Aperture
Grid

## Slide 14: Centroid Measurements: Beam Position and Phase
Centroid is primarily measured by the Beam Position Monitor (BPM) system
  In most accelerators, transverse position is the most important measurement
  In a linac like at ESS, longitudinal position (phase) is also very important
Additional systems at ESS
  Profile monitor measures position after ion source (no appropriate signal for BPMs)
  Imaging and grid systems measure centroid on target, dump
  Bunch Shape Monitor also reports phase centroid vs. RF reference

## Slide 15: Centroid Measurement: Measure Position of Beam with Electromagnetic Beam Position Monitors
Simplified model: image charge in an electrode is proportional to the beam current and to the angle included between the beam and the electrode extremes:
In addition to this geometric effect, the field lines tend to cluster closely in the region of the nearest electrode (the E field must be perpendicular to the walls). For this geometry, this gives an additional factor two.
J. Byrd
Pro Tip: Analytical approximations for concept development, FEM for detailed design

## Slide 16: Centroid Measurement: Position monitor styles They also measure time of beam arrival
SOR-Ring
Buttons
Striplines
Byrd, Musson

## Slide 17: Centroid Measurements: Beam Position and Phase Electronics
  Vector receiver
  352 MHz and 704 MHz
  non-IQ receiver design coordinated with Radio Frequency (RF) team
Approximately 1 mA of beam
Approximately 0.1 mA of beam
S. Molloy, R. Baron, ESS Lund
Analog transition module, Digitizer with FPGA
Also, custom RF front end

## Slide 18: Distribution Measurements
Peer into the the beam and measure with more detail
  Transverse Phase Space
  Transverse Profile
  2-dimensional map of transverse distribution
  Longitudinal Profile (of the short microbunches)
Systems
  Emittance Measurement Units
  Wire Scanners
  Bunch Shape Monitors
  Non-invasive Profile Monitors
  Grids
  Imaging Systems

## Slide 19: Distribution Measurement: Emittance measurement of beam at low energy
Allison scanner developed at CEA for the ESS LEBT. Full power capable.
Schematic view of Allison scanner (CEA/IRFU)
LINAC4 emittance meter slit and grid (CERN)
At 75 keV
At 3.6 MeV

## Slide 20: Distribution Measurement: Measure Transverse Profile of Short Beam Pulses with Wire Scanners
DTL example shown Wires located all along the accelerator
Current induced in wire vs. wire position
Temperature of wire vs. wire position
B. Cheymol

## Slide 21: Distribution Measurements: Invasive Profile and Shape
analysis: B. Cheymol
Linear wire scanners (as previous slide) and Bunch Shape Monitors
Fast CERN wire scanners in High energy transport line
Multi-wire grid in target station
Like a wire scanner, but a grid of stationary wires
Increasing beam handling capability

## Slide 22: Distribution Measurement: Measure Long Beam Pulses with Non-invasive Profile Monitors
Gas Luminescence monitor
Very little gas in cold part of accelerator, so there, we use more efficient, but complicated method of collecting ions inside the vacuum chamber
C. Thomas

## Slide 23: Distribution Measurement: Measure 2-D Map of Beam Current with Imaging Systems
Luminescent coating on target provides image; transported through complex optics to camera
Mockup of Desired Beam-on-Target Display

## Slide 24: Distribution Measurement: ESS Imaging Systems Technologies
Retroreflector for the fiducials
Optics
Development
Ray Tracing results
Target Design (simpler for dump)
Prototype @Oslo
Luminescent Material Development
Spectrum
Yield
Decay time
Temperature and Radiation Tolerance
Temperature
Wavelength
Al2O3:Cr
Irradiation tests:
Y2WO6:Sm
Spectrum, Yield
Capsule @BNL
Also @DTU
Several alternatives
Baseline
Additional Development
Beam tests: @ORNL, LANL, Oslo, CERN
XRD, XANES / XAS,TEM
Coating technology: HV, Stony Brook

## Slide 25: Protection: from Beam Loss
  Total beam loss, microsecond measurement latency required for protection
  BCM, icBLM (saturation, nBLM (current mode) -> Interlock; Threshold/derivative term for fast protection
> 1.6 milliamp lost for up to 200 µs
  BCM, icBLM, nBLM -> Interlock; Damage model for protection
~ microCoulomb lost over 200 µs to many seconds (diffusion time)
  icBLM, nBLM -> Interlock; Damage model for protection
~ “1 Watt/meter” radiation dose management
  icBLM, nBLM -> alarm based on dose/activation plan
Sparse coverage by Current Monitors
coverage by Ion Chambers
Neutron Detectors

## Slide 26: Protection: from Damaging Beam Conditions @Target/Dump
Candidate interlock inputs
BPM (target)
Aperture
Grid
Imaging
By By Bx Bx | Bx Bx By By
Quads 1-4
Quads 5,6
Beam on Target
Raster magnets
Magenta line: beam centroid Blue line: 10 rms
Neutron Shield Wall
2 m
4 cm

## Slide 27: ESS Diagnostics Partners and Collaborators
Aarhus University
CEA Saclay, Paris
CERN, Geneva
Cockcroft Institute, Daresbury
DESY, Hamburg
Elettra – Sincrotrone Trieste
ESS Bilbao
INFN, Catania
INFN, Legnaro
Lund University
University of Oslo
Technical University of Denmark
Science and Technology Facilities Council, Daresbury
Warsaw University of Technology
Chinese ADS
J-PARC, Japan
Oak Ridge National Laboratory
Los Alamos National Laboratory
INR, Moscow
Högskola Väst, Trollhättan

## Slide 29: Example from SNS in the US: First Beam on Target
ESS in Sweden: First beam next year, beam on target in a few short years

## Slide 30: Thank you