## Slide 1: icBLM electronics
Clement Derrez
Test Engineer
www.europeanspallationsource.se
15 November, 2017

## Slide 2: Outline
Conceptual system design
Detectors
Electronics
  High voltage unit
  Signal acquisition unit
Conclusion & Outlook

## Slide 3: Conceptual system design
BLM system: One of the most important diagnostic system of the ESS LINAC.
  Keep the machine safe from beam-induced damage
  Avoid excessive machine activation by providing critical input to the MPS
   no blind spots & precise loss locations determination
System designed for maximum reliability
  Redundancy
  low latency
Challenges:
  Complicated noise situation. Additionally, noise sources vary depending on the location within the accelerator.
  Cabling: new requirement gives us cables length up to 100m

## Slide 4: Conceptual system design
Detectors: ICs
HV
Signal acquisition
FPGA processing,
interfacing
EPICS IOC
Timing receiver
User screens,
Archiver…
microTCA +/ custom CERN crate
ESS control room
Tunnel
Gallery
MPS
Post mortem
Timing

## Slide 5: Detectors
  LHC-type
  Design criteria: signal speed & robustness against ageing
  Operated at 1.5 kV
  Low pass filter at HV input
  285 ionization chambers received at ESS in Jul. 2017

## Slide 6: Electronics: Signal acquisition unit specifications
Signal acquisition: Dynamic range | Input current dynamic range: 10 nA to 10 mA. | Two separate ranges shall be used: the low end set by the condition that the operators must be able to tune on beam loss measurements down to 10 mW/m, while the high end shall allow recording of total beam loss during faults and machine studies | BLM is required to be able to measure at least 1% of 1W/m loss during normal operation and up to 1% of the total beam loss.
Signal acquisition: Response time | NC LINAC: Calculated melting time values of 3-4 us SC LINAC: 10 us The signal acquisition card shall therefore feature a bandwidth of at least 300 KHz and a minimum sampling rate of 1MSPS | “Particle time” - PT: time between the onset of beam loss (the primary is lost) and the moment particle (primary or secondary) hits the detector. Detection time: time needed for the detector signal to develop and to collect enough hits/current Processing time: from the output of the detector to the BIS output on the FPGA | 
Signal acquisition: resolution | The signal acquisition card shall have a resolution of at least 20 bits |  | 
Signal acquisition: noise figure | The signal acquisition card shall have a typical noise better than 10 nA in its lowest range, including temperature induced noise |  | 
 We need a 300 KHz Bandwidth
 We want to be able to trigger on a 800 nA, 5 us pulse width, with 100m max cable length
Notes: The response time of the BLM system is a sum of:
  “Particle time” - PT: time between the onset of beam loss (the primary is lost) and the moment particle (primary or secondary) hits the detector.
  Detection time: time needed for the detector signal to develop and to collect enough hits/current
  Processing time: from the output of the detector to the BIS output on the FPGA
In a word: Dynamic range specs:
  “BLM is required to be able to measure at least 1% of 1W/m loss during normal. operation and up to 1% of the total beam loss”.
  Gave estimation on the ICLBM current range: ~800nA – few mA

## Slide 7: Electronics: Data acquisition, 2 options
COTS solution: CAENels PICO-8
  8-channel 20 bit Bipolar Floating Picoammeter with MTCA.4 REAR I/O
  full-scale ranges of 10 mA and 500 uA
  input bandwidth: 300 kHz
  “manual” change between ranges
BLEDP: Current to Frequency Converter, fully differential integrator circuit.
  Current to frequency converter, fully differential front end
  Measurement of the current input is performed by two different techniques:
    Advanced Current to Frequency Converter (ACFC) used in the range 10pA to 30mA
    Direct ADC acquisition (DADC) used in the range 20.3µA to 200mA
    Transition between ranges managed by an FPGA
    Optical output to post processing unit
    Bandwidth: 1 MHz

## Slide 8: Electronics: Data acquisition
Noise levels, 250 KHz sampling rate (fits BLEDP):
  Pico8 range: 500 uA
   We need to find a lower capacitance cable if we go with the pico8
 | PICO8 | BLEDP
100m triax ( ~100pF/m) | 291 nA | 130nA
 | PICO8 | BLEDP
100m coax ( ~50 pF/m) | 190 nA | 130nA
 | PICO8 | ESS testbench
2m triax | 10 nA | 100 nA

## Slide 9: Electronics: Data acquisition, 2 options

## Slide 10: Debugging still needed on BLEDP
Spikes in data out at current levels > 10 uA
BLEDP only measures positive current.

## Slide 11: Some more points to consider
BLEDP
  Resources for support
  Custom crate design
  Reliability analysis available (10 years MTTF)
PICO8
  Range needs to be changed “manually”

## Slide 12: Electronics: HV unit
COTS solution: CAENels HV-PANDA
  Double-width AMC with MTCA.4 Rear I/O
  4 channels rated at 4kV at 7W
  Configurable ramp rate
  Nominal voltage accuracy better than 0.05% FS and a ppm-level peak-to-peak output voltage ripple
  SHV type connectors

## Slide 13: Electronics: HV unit
optimal freq: 0.2 Hz

## Slide 14: Conclusions & Outlook
System design is moving forward quickly. 285 ICs received last week
FRU tests:
  Perform a “long-term” measurement to check for possible drift, on HV and signal acquisition units
Software and Firmware