## Slide 1
Multi-Blade 10B-based detector for
neutron reflectometry at ESS:
tests at the
CRISP reflectometer at ISIS - RAL (UK)
Francesco Piscitelli
on behalf of
ESS / Wigner / Lund University / Linköping University collaboration
IEEE NSS Atlanta
2017/10/26
Wigner Research Institute
Building research infrastructure and synergies for highest scientific impact on ESS
Linköping University

## Slide 3
ESS will be several times brighter than existing facilities
Linear proton accelerator
target
instruments
August 2015
Powerful neutron instruments
Sweden

## Slide 4
ESS will be several times brighter than existing facilities
Powerful neutron instruments
Linear proton accelerator
target
2 Reflectometers:
FREIA: horizontal refl.
ESTIA: vertical refl.
instruments
August 2015

## Slide 5
ESS will be several times brighter than existing facilities
Powerful neutron instruments
High detector requirements
1 10 100
1 = state-of-the-art
Instantaneous
Rate Capability
100 kHz/mm2
1 2 3 4
Reflectometry
Spatial Resolution
Reflectometry
0.5 mm
Linear proton accelerator
target
2 Reflectometers:
FREIA: horizontal refl.
ESTIA: vertical refl.
instruments
August 2015

## Slide 6
The Multi-Blade project
10B-detector for reflectometers

## Slide 7
The Multi-Blade project
Wigner Research Institute
Lund
Budapest
The key objective of WP4 is the technological
evolution of neutron detectors in terms of
resolution, intensity and dimensions.
3 years
Task 4.2 Neutron Detectors – The Intensity Frontier
Linköping University
Linköping
10B-detector for reflectometers

## Slide 8
The Multi-Blade project
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
A cassette (unit)
High counting rate capability
High spatial resolution
10B-detector for reflectometers

## Slide 9
The Multi-Blade project
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
A single Boron layer inclined at 5 degrees
A cassette (unit)
neutrons
θ
neutrons
θ
10B4C
10B4C
q = 5 degrees
q = 90 degrees
Efficiency 45% at 2.5Å
Efficiency <5% at 2.5Å
High counting rate capability
High spatial resolution
10B-detector for reflectometers

## Slide 10
The Multi-Blade project
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
Why the counting rate capability is improved?
The intensity is spread over a wider surface (5 degrees ~ factor x10)
Thin gap MWPC (4mm)
Low gas gain operation G~20
(max 0.2pC avalanches)
A cassette (unit)
High counting rate capability
High spatial resolution
10B-detector for reflectometers

## Slide 11
The Multi-Blade project
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
cathode strips
anode wires
holding frame
10B4C substrate
A cassette (unit)
10B-detector for reflectometers

## Slide 12
The Multi-Blade project
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
cathode strips
anode wires
holding frame
10B4C substrate
Ti blades
32 Strips
32 Wires
A cassette (unit)
A cassette (unit)
10B-detector for reflectometers
A cassette (unit)

## Slide 13
The Multi-Blade project
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
neutrons
neutrons
130mm
100mm
A cassette (unit)
32 wires
(front)
32 Strips
(back)
9 units

## Slide 14
The Multi-Blade project
neutrons
neutrons
130mm
100mm
32 Strips
(back)
9 units
32 wires
(front)
vessel

## Slide 15
The Multi-Blade project
vessel
Front-end on wires
9 boards (32x9ch)
Front-end on strips
9 boards (32x9ch)
32ch front-end boards

## Slide 16
The Multi-Blade project
Front-end on wires
9 boards (32x9ch)
Front-end on strips
9 boards (32x9ch)
cassette
32 wires
32 strips
x9
FEE
vessel
FEE
x9
x9
32x9 outputs
from FEE strips
Each blade has 32 wires and 32 strips
readout individually,
in total 9 units (64x9 channels)
CAEN
V1740D
One Digitser (64ch is 1 cassette)
TTL for time stamp reset

## Slide 17
Summary …before the tests
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).
 | Multi-Blade
efficiency | 45% @ 2.5Å 56% @ 4.2Å 65% @ 5.1Å
spatial resolution | 0.5 x 2.5 mm2
uniformity | 10%
stability | <1% in 12h
counting rate capability | >1.6 kHz/mm2 (lower limit) >17kHz / channel (lower limit)
gamma-ray sensitivity | < 10-7 (with 100keV threshold)
gas gain | 20
overlap | 50% eff. drop in 0.5mm gap

## Slide 18
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).
 | Multi-Blade
efficiency | 45% @ 2.5Å 56% @ 4.2Å 65% @ 5.1Å
spatial resolution | 0.5 x 2.5 mm2
uniformity | 10%
stability | <1% in 12h
counting rate capability | >1.6 kHz/mm2 (lower limit) >17kHz / channel (lower limit)
gamma-ray sensitivity | < 10-7 (with 100keV threshold)
gas gain | 20
overlap | 50% eff. drop in 0.5mm gap
Matching ESS requirements
Summary …before the tests

## Slide 19
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).
 | Multi-Blade
efficiency | 45% @ 2.5Å 56% @ 4.2Å 65% @ 5.1Å
spatial resolution | 0.5 x 2.5 mm2
uniformity | 10%
stability | <1% in 12h
counting rate capability | >1.6 kHz/mm2 (lower limit) >17kHz / channel (lower limit)
gamma-ray sensitivity | < 10-7 (with 100keV threshold)
gas gain | 20
overlap | 50% eff. drop in 0.5mm gap
Matching ESS requirements
Summary …before the tests

## Slide 20
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).
 | Multi-Blade
efficiency | 45% @ 2.5Å 56% @ 4.2Å 65% @ 5.1Å
spatial resolution | 0.5 x 2.5 mm2
uniformity | 10%
stability | <1% in 12h
counting rate capability | >1.6 kHz/mm2 (lower limit) >17kHz / channel (lower limit)
gamma-ray sensitivity | < 10-7 (with 100keV threshold)
gas gain | 20
overlap | 50% eff. drop in 0.5mm gap
Matching ESS requirements
x3 better than
state-of-the-art
x10 better than
state-of-the-art
Summary …before the tests

## Slide 21
ToF Reflectometry
2 Reflectometers @ ESS:
FREIA: horizontal refl.
ESTIA: vertical refl.
10B-detector for reflectometers

## Slide 22
ToF Reflectometry
detector
sample
2 Reflectometers @ ESS:
FREIA: horizontal refl.
ESTIA: vertical refl.
10B-detector for reflectometers
Reflectometry is a technique to study
SURFACES & INTERFACES

## Slide 23
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
ToF Reflectometry
detector area
reflected intensity

## Slide 24
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
ToF Reflectometry
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)

## Slide 25
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
ToF Reflectometry
detector area
reflected intensity
0.5 6

## Slide 26
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
ToF Reflectometry
detector area
reflected intensity
0.5 6
l1
l2
l3

## Slide 27
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
ToF Reflectometry
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
l1
l2
l3
v1
v1
v1
Faster neutrons
Slower neutrons
Larger l
Smaller l
l∝1/v = ToF / Distance
(0.5Å)
(6Å)
0.5 2 6Å
l1
l2
l3
0.5 6
T1
Smaller ToF
Larger ToF
T2
T3

## Slide 28
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
ToF Reflectometry
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
l3
l1
l2
q =(4p/l) sin(q)
Incidence angle
Neutron wavelength
Log R =
= reflected
/ incoming
q
l∝1/v = ToF / Distance
l1
l2
l3
0.5 2 6Å
0.5 6
Line at constant q
q ∝sin(q) / l
Momentum transfer

## Slide 29
CRISP reflectometer at ISIS
Tests performed 2 weeks ago on the CRISP reflectometer at ISIS

## Slide 30
CRISP reflectometer at ISIS
Thermal Neutron Reflectometer | 
Neutron wavelength | 0.5 – 6.5 Å @ 50Hz (0.5 – 13 Å @ 25Hz)
q range | 0.005 – 1.1 Å-1
Moderator - sample distance | 10.25m
Sample - Detector distance | 1.87m
Sample - MB Detector distance | 2.33m
detector
sample

## Slide 31
CRISP reflectometer at ISIS
Thermal Neutron Reflectometer | 
Neutron wavelength | 0.5 – 6.5 Å @ 50Hz (0.5 – 13 Å @ 25Hz)
q range | 0.005 – 1.1 Å-1
Moderator - sample distance | 10.25m
Sample - Detector distance | 1.87m
Sample - MB Detector distance | 2.33m
detector
sample
detector
sample

## Slide 32
sample
MB detector
sample
CRISP reflectometer at ISIS
Multi-Blade
detector

## Slide 33
sample
MB detector
sample
CRISP reflectometer at ISIS
Multi-Blade
detector
130mm
100mm
The cassette (units)
are placed horizontally
Resolution 0.5mm
Resolution 4mm

## Slide 34
Results

## Slide 35
40mm
60mm
sample
Results
BN mask
Multi-Blade
detector

## Slide 36
40mm
60mm
sample
Results
BN mask
Multi-Blade
detector
(Log scale)

## Slide 37
40mm
60mm
sample
Results
BN mask
shadow
Multi-Blade
detector
(Log scale)

## Slide 38
strips
wires
40mm
60mm
shadow
sample
Results
BN mask
10B4C
substrate
shadow
Multi-Blade
detector
(Log scale)

## Slide 39
strips
wires
40mm
60mm
shadow
sample
Results
BN mask
Not a dead area but the overlap avoids dead zones
10B4C
substrate
shadow
shadow
No shadow
Multi-Blade
detector
(Log scale)

## Slide 40
BN mask
Results
60mm
1.5mm
(Log scale)
130mm

## Slide 41
Multi-Blade
detector
sample
Results
Ir sample
Fe/Si sample
Si sample
q
q
Fe/Si
Ir
Si

## Slide 42
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
ToF Reflectometry
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
l3
l1
l2
q =(4p/l) sin(q)
Incidence angle
Neutron wavelength
Log R =
= reflected
/ incoming
q
l∝1/v = ToF / Distance
l1
l2
l3
0.5 6
0.5 2 6Å
Line at constant q
q ∝sin(q) / l
Momentum transfer

## Slide 43
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
0.5 2 6Å
0.5 6
Ir sample
Multi-Blade
detector
(Theoretical R)

## Slide 44
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
q1
q2
q2
q1
0.5 2 6Å
0.5 6
Multi-Blade
detector

## Slide 45
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
q =(4p/l) sin(q)
l3
l1
l2
Log R =
= reflected
/ incoming
q
l3
l1
l2
q1
q2
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
q1
q2
q2
q1
0.5 6
0.5 2 6Å
Multi-Blade
detector

## Slide 46
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
q1 = 0.2deg
q2
q1
q7 = 0.8deg
…
0.5 2 6Å
0.5 6
Multi-Blade
detector

## Slide 47
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
q1 = 0.2deg
q2
q1
q7 = 0.8deg
…
Background
0.5 6
0.5 2 6Å
Multi-Blade
detector

## Slide 48
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
q2
q1
Background
Background
0.5 2 6Å
Multi-Blade
detector
(Log scale)

## Slide 49
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
q2
q1
Background
Background
Gating between 6ms and 20ms
0.5 2 6Å
Multi-Blade
detector
(Log scale)
(Log scale)

## Slide 50
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Results
Background
Background
Gating between 6ms and 20ms
100 neutrons
strips Y
wires X
substrate
10B4C
2Å
90%
98%
61%
45%
Al Ti SS
Nominal 10B4C coating 7um
Actual 10B4C coating 4um
Blades
0.5 2 6Å
(Log scale)
(Log scale)
55 neutrons get scattered @ 0.5Å
0.5Å
MB is optimized for cold wavelengths
and
CRISP is a thermal reflectometer

## Slide 51
Results
6ms (2Å)
20ms (6Å)
any ToF
ToF > 6ms (2Å)
Background identified and understood

## Slide 52
Results
6ms (2Å)
20ms (6Å)
ToF > 6ms (2Å)
All curves match for all angles and neutron wavelengths
The MB uniformity is good enough
Line at constant q
q1 = 0.2deg
q7 = 0.8deg
…
0.5 6
Line at constant q
l1
l4

## Slide 53
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
0.5 2 6Å
0.5 6
Multi-Blade
detector

## Slide 54
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
Conventional
Conventional
0.5 6
0.5 2 6Å
q
Multi-Blade
detector

## Slide 55
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
Conventional
Incl. Res.
Incl. Res.
Conventional
0.5 6
0.5 2 6Å
Integration at constant q,
correcting q with the position
on the detector
q
Multi-Blade
detector

## Slide 56
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
Conventional
Incl. Res.
Incl. Res.
Conventional
The 0.5mm spatial resolution is needed to calculate the right q
0.5 6
0.5 2 6Å
Constant q
q + Dq
Same l
q - Dq
q1 =(4p/l) sin(q+Dq)
q2 =(4p/l) sin(q-Dq)
Integration at constant q,
correcting q with the position
on the detector
q
q
Multi-Blade
detector

## Slide 57
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
Conventional
Incl. Res.
Incl. Res.
Conventional
The 0.5mm spatial resolution is needed to calculate the right q
Conventional
Incl. Resolution
0.5 6
0.5 2 6Å
Constant q
q + Dq
Same l
q - Dq
q1 =(4p/l) sin(q+Dq)
q2 =(4p/l) sin(q-Dq)
Integration at constant q,
correcting q with the position
on the detector
q
q
Multi-Blade
detector

## Slide 58
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
sample
q
q
COLLIMATED
DIVERGENT
HIGHER
RATE
LOWER
RATE
0.5 6
0.5 6
Multi-Blade
detector
Multi-Blade
detector

## Slide 59
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
sample
q
q
COLLIMATED
DIVERGENT
HIGHER
RATE
LOWER
RATE
0.5 6
0.5 6
Integration at constant q,
correcting q with the position
on the detector is mandatory
Multi-Blade
detector
Multi-Blade
detector

## Slide 60
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
COLLIMATED (120 min)
sample
q
q
COLLIMATED
DIVERGENT
DIVERGENT (14 min)
The MB 0.5mm resolution and rate capability is needed to measure in divergent mode
HIGHER
RATE
LOWER
RATE
0.5 6
0.5 6
Integration at constant q,
correcting q with the position
on the detector is mandatory
Multi-Blade
detector
Multi-Blade
detector

## Slide 61
sample
Results
q
q
Ir sample
Fe/Si sample
Si sample
DIVERGENT
>5 orders of magnitude with a poor shielding for background
100
10-5
0.5 6
q - Dynamic range
Multi-Blade
detector
(Usually 10-6 is obtained in experiments)

## Slide 62
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
Multi-Blade
detector

## Slide 63
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
0.5 2 6Å
Multi-Blade
detector

## Slide 64
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
ToF (or l)
detector area
detector ToF
0.5 2 6Å
Multi-Blade
detector

## Slide 65
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
ToF (or l)
Specular
Specular
Off-Specular
Off-Specular
0.5 2 6Å
Multi-Blade
detector

## Slide 66
ToF (or l)
Specular
Specular
Off-Specular
Off-Specular
0.5 2 6Å
Ir sample
Fe/Si sample
Si sample
sample
Results
q
q
Multi-Blade
detector

## Slide 67
Results
Specular Reflectivity
Supermirror edge
(m=3.8)
Si edge
Correlated roughness
domains from the layers
Spin-flip scattering
signal from the layers
Sample Horizon
Transmitted beam
through the sample
No Transmitted beam,
total reflection
Off-specular scattering from Fe/Si neutron supermirror

## Slide 68
Conclusions
Specular Reflectivity
Supermirror edge
(m=3.8)
Si edge
Correlated roughness
domains from the layers
Spin-flip scattering
signal from the layers
Sample Horizon
Transmitted beam
through the sample
No Transmitted beam,
total reflection
Off-specular scattering from Fe/Si neutron supermirror
All the features of the samples have been clearly identified

## Slide 69
Conclusions
Specular Reflectivity
Supermirror edge
(m=3.8)
Si edge
Correlated roughness
domains from the layers
Spin-flip scattering
signal from the layers
Sample Horizon
Transmitted beam
through the sample
No Transmitted beam,
total reflection
Off-specular scattering from Fe/Si neutron supermirror
All the features of the samples have been clearly identified
Science can profit from the features of this detector

## Slide 70
Conclusions
Specular Reflectivity
Supermirror edge
(m=3.8)
Si edge
Correlated roughness
domains from the layers
Spin-flip scattering
signal from the layers
Sample Horizon
Transmitted beam
through the sample
No Transmitted beam,
total reflection
Off-specular scattering from Fe/Si neutron supermirror
All the features of the samples have been clearly identified
Science can profit from the features of this detector
The Multi-Blade detector can be considered
a mature technology for neutron reflectometry
… now it is time for the fine tuning ...

## Slide 71
Thank you

## Slide 72
Backup slides

## Slide 73
The Multi-Blade project
Front-end on wires
9 boards (32x9ch)
Front-end on strips
9 boards (32x9ch)
cassette
32 wires
32 strips
x9
FEE
vessel
FEE
x9
x9
32x9 outputs
from FEE strips
CAEN
V1740D
One Digitser (64ch is 1 cassette)
TTL for time stamp reset
Strip n
Strip n+1
Wire m
Wire m+1
Strip n-1
Strip n
Wire m
Strip n+1
Example of clusters per neutron event
Each blade has 32 wires and 32 strips
readout individually,
in total 9 units (64x9 channels)

## Slide 74
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Ir sample
Fe/Si sample
Si sample
Multi-Blade
detector
sample
Results
q
q
q2
q1
Background
Background
Gating between 6ms and 20ms
10-3
10-4

## Slide 75
Edge
gain not corrected
Garfield Simulation
Measured
Edge
wires
Uniformity and overlap
gain corrected

## Slide 76
Efficiency
64% @ 5.1Å
measured
calculated
58% @ 4.2Å
45% @ 2.5Å

## Slide 77
PHS and Gain
Gas Gain ~ 20 (max 0.2pC avalanche charge)
(contributes to improve the counting rate capability)
Energy Resolution ~50 KeV
Energy Threshold ~100 KeV
Measured @ 4.1Å 5 degrees
measured
calculated*
*F. Piscitelli and P. van Esch,
Analytical modeling of thin film neutron converters and its application to thermal neutron gas detectors
J. Instrum. 8, P04020 (2013)
arXiv:1302.3153

## Slide 78
Spot
Counting rate capability
No saturation observed up to ~ 1.6kHz/mm2
Slit
Rate per mm2
Rate per channel
17kHz
1.6kHz/mm2
Beam flux (a.u.)

## Slide 79
Spatial resolution X – wires
0.35mm
Neutron beam
(Step 0.25mm)
wires
Spatial Resolution
equivalent FWHM* ~ 0.6mm
*P. Van Esch et al.,
An information-theoretical approach to image resolution applied to neutron imaging detectors based upon discriminator signals
in Proceedings of ANNIMA conference, Marseille 2013
arXiv:1307.7507

## Slide 80
strips
wires
shadow
sample
Results
BN mask
Not a dead area but the overlap avoids dead zones
10B4C
substrate
shadow
No shadow
dead area
No shadow = No overlap
Shadow = Overlap
Multi-Blade
detector