## Slide 1: Status of MultiGrid Detector Development for CSPEC and TREX
Anton Khaplanov
Isaak Lopez Higuera
www.europeanspallationsource.se
Sep 28, 2017

## Slide 2: Outline
B10 detector development
Multi-Grid detector Characterization highlights
  Intrinsic background suppression
  Gamma sensitivity
  Current development
Multi-Grid ToF results from CNCS, SNS
Engineering the Multi-Grid
MG.CNCS
MG.SEQ
Conceptual design for CSPEC

## Slide 3: B10 Detectors as Alternative to He3
Large area
Low cost
Modular
High total efficiency
Many layers – lower local rate
No need for resistive readout – low gain
B-10 film
on both sides
Al substrate
n
conversion
gas
fragment 1
fragment 2
gas
Anode wires
n
fragment 1
fragment 2
Multi-Grid:
Multiple layers at normal incidence
~5% efficiency
in 1 layer
n
n
Introduced at ILL, jointly developed by ILL and ESS under CRISP project,
now under BrightnESS

## Slide 4: Multi-Grid Prototypes for Characterization
Measurements:
Absolute efficiency
Uniformity
Position reconstruction
γ sensitivity
Rate tests
Ageing
MG.24 @R2D2 @IFE
Proto 2 at CT2 @ILL
Proto 1 at CT1 @ILL

## Slide 5: Multi-Grid How it works: Grid
Layer thicknesses for cold spectrometer optimization used at MG.CNCS and applies to CSPEC(16 blades):
7 blades 0.5μm,
7 blades 1.0μm,
3 blades 1.5μm
For thermal optimization 20 blades total:
4 blades 1.0μm,
10 blades 1.25μm,
6 blades 2.0μm.
Column assembly
Grid
n
Notes: In the Multigrid, rather than a casette we have grids, each grid has 17 voxel per column and 4 columns. Each voxel is 22.5x22x11mm. In case we aim for thermal neutrons, we would have 21 voxels. The vertical blade of each voxel is coated, so the neutrons travel perpendicularly to that blade. The layer thickness is optimized for the position in the grid, so not all blades have the same thickness, this allows us to increase effectivity. In the middle of every voxel we have the wire (composition of the wire, microns, resistive?). So when we have the signal, we will know exactly in which position we are with the 2 coordinates, wire and voxel, X and Y.

## Slide 6: Multi-Grid How it works: column assembly
n
Notes: All the grids are mounted onto a bottom insulator, usually made in PEEK, and these onto a PCB that reads the grid signals. On both sides of the grids, we have the end-grids. These hold the PCBs with the pads where we solder the wires. All this is mounted into the backplate, where we also have the HV input, gas inlet and outlet and the readout feedthroughs.
Explain the readout models.

## Slide 7: First Multi-Grid tested on an Instrument
Detector tested on cold chopper spectrometer IN6 @ILL
Active Area 30 x 45 cm2
96 grids in 6 modules of 16 grids, 360 anode wires
Operated for ~2 weeks, replacing 25 He3 tubes
@ 4.1, 4.6 and 5.1 Å incoming wavelength
Main topics of the test:
First ToF spectra
Background investigation
Elastic line shape
*A. Khaplanov et al. 2015 JINST 10 P10019
*A. Khaplanov et al. J. Phys: Conf Ser 528 (2014) 012040

## Slide 8: Intrinsic background
Spectrometers work with a very large dynamic range  very high signal-to-noise is required.
Intrinsic background often found in Al due to α emitter impurities (U and Th).
Spectrum from standard Al
N + 10B7Li + α
7Li α
α decay U, Th
x50-x100 background reduction using:
high-purity Al
Ni-plating
Ni-plating on a 0.5mm Al blade
A. Khaplanov et al. 2015 JINST 10 P10019

## Slide 9: γ Sensitivity Investigated
* A. Khaplanov et al., JINST 8, P10025 (2013), arXiv:1306.6247
Operation HV, can choose γ sensitivity 10-6 – 10-9
High statistics 137Cs measurement
Plateau for γ and n measurement
Beam + source setup for γ sensitivity measurement
γ rejection demonstrated in ToF @IN6

## Slide 10: Test on CNCS, Cold Neutron Chopper Spectrometer
A kind offer by K. Herwig to test MG at SNS
Recommendation of 2015 ESS annual review
Goals:
Test at spectrometer
Side-by-side comparison to He3
In parallel with user experiments
Dedicated tests
Solution:
Size = half of “8-pack” module – 1.1m x 19cm
Installation June-July 2016
Operated July 2016 to June 2017

## Slide 11: Multi-Grid installed at CNCS
Detector installed at 57° scattering angle
Inaccessible between summer and winter shutdowns
DAQ setup outside, accessible remotely
MG
He-tubes
3He 8-packs
M-G
1.1m
19cm
20cm
View from sample
2m
Front-end
electronics, Mesytec
DAQ
outside detector tank

## Slide 12: Methods Used for Tests
Vanadium
Counts over time, background
Fast neutrons (fission, prompt)
Water, Ice
Allowing to measure:
Energy resolution – efficiency, scattering – background sensitivity – saturation
Gamma background
(NaI detector)
x2 lower fast n counts in MG
Long term operation
Single crystal
High local rate

## Slide 13: Elastic Energy Resolution

## Slide 14: Efficiency
Counts from integral of the elastic peak in dE
Efficiency measured relative to the nearest He3 tubes
Error bars: 5-8%, agrees with prediction

## Slide 15: Single Crystal Reflection
Crystal
Incident beam
Reflected beam at Ei = 17.20meV
Reflected beam at Ei = 13.74meV
A single crystal reflects neutrons according to Bragg’s law:
Resulting in an intense spot seen by detector
He3
MG
He3
MG
No loss of position resolution or saturation observed in Multi-Grid

## Slide 16: Energy Ranges for ESS Spectrometers
Energy ranges of the first ESS ToF spectrometers
Cold range, 0.76 to 15 meV measured
Measurements needed for thermal range, up to 160 meV
 | CSPEC | T-REX
Typical initial λ, Å (meV) | 2 to 15 Å (20 to 0.36 meV) | 0.7 to 6.4 Å (160 to 2 meV)
Incident energies used at CNCS and SEQUOIA (M. Stone et al.)

## Slide 17: Efficiency Optimizations
Optimized grid
1.5μm
0.5μm
1.0μm
n
Layer thicknesses in MG.CNCS (16 blades):
7 blades 0.5μm,
7 blades 1.0μm,
3 blades 1.5μm
Efficiency as a function of wavelength for MG optimized for cold vs. thermal spectrum
Cold – optimization centered on 4Å
Thermal – optimization centered on 1Å
Thermal optimization
Thermal Optimization
Tested at MG.CNCS
*F. Piscitelli, P. van Esch, 2013 JINST 8, P04020
20 blades total:
4 blades 1.0μm,
10 blades 1.25μm,
6 blades 2.0μm.

## Slide 18: Multi-Grid Thermal Demonstrator test
Proposal made to do the test at SEQUOIA, at SNS
In addition to what was done for CNCS test:
Performance in thermal (and epithermal) flux
Efficiency optimization
Energy resolution
Line shape
Scattering
Detector construction
Vacuum compatibility
Online data acquisition, visualization and diagnostics
(to be developed by DMSC)

## Slide 19: Designing and Building Multi-Grid
ESS Detector Group Engineers:
Michail Anastosopoulos
Isaak Lopez Higuera
Ioannis Apostolidis
Embla facility in Lund
Utgård facility in Lund

## Slide 20: Multi-Grid demonstrator MG24 and iterations
MG24: First demonstrator
MG24.Copper: Test with low α
MG24.T: Optimized for thermal neutrons
MG24.Alu foil: Test with foil window
Notes: This is the first test demonstrator we built, it was 2 columns of 12 grids each. The window thickness was 2mm. We tested it for XXX
We did the same test but with copper blades.
And then we did the same but for thermal neutrons, so 21 voxels.
Also, we have another vessel with interchangeable windows, so we can put aluminum foil.
Explain the basics of MG24, columns, gaps etc
The needs of MG.Copper and thermal versions
Tests done
Need for aluminum foil...probably not tested yet

## Slide 21: MG.CNCS
Grid assembly
End-Grid
Backbone
Bottom PCB
Bottom insulator

## Slide 22: MG.CNCS
Installed at SNS in the CNCS instrument during June 2016
Optimized for cold neutrons
2 columns of 48 grid each
Active area 1100 x 185 mm
Voxel volume 22.5 x 22 x 11 mm
1394 x 200 x 250mm

## Slide 23: MG.SEQ Installation proposal at SNS in the SEQUOIA instrument
Test with thermal/epithermal neutrons
First large area detector under vacuum
Mounted in the lower bank ~5m and 25˚ from sample

## Slide 24: Multi-Grid demonstrator MG.SEQ: Installation proposal at SNS in the SEQUOIA instrument
9 columns, 40 grids in each
Telescopic sliders to access grid assemblies
Use of inclined wedges so all grids are at the same distance from the sample
1050 x 914 x 450mm
neutrons
MG.SEQ isometric view
MG.SEQ top view

## Slide 25: Multi-Grid demonstrator MG.SEQ: Installation proposal at SNS in the SEQUOIA instrument
3mm thickness window
8 rows of 17 rods, Ø3mm
Full access only from 1 side
n
MG.SEQ bottom view
Total deformation

## Slide 26: Multi-Grid demonstrator MG.SEQ: Installation proposal at SNS in the SEQUOIA instrument
Amplified 50 times
Max deformation: 1.2mm on the sides
Max equivalent stress: 176 MPa
Stress in the rods: 80 MPa
Equivalent stress in the vessel
Equivalent stress in the rods
Amplified 50 times

## Slide 27: CSPEC at ESS
Cold spectrometer
0.2 meV < Ei < 20 meV
29m2 detector
Horizontal coverage 5° to 135°
Vertical coverage -25° to 25°

## Slide 28: Multi-Grid Detector for CSPEC
Neighbour vessels configuration, minimizing dead spaces, 21mm from grid to grid
Integration of the detector vessels in the CSPEC instrument
Vessel length: 3.5m
Detector configuration: 2 columns per vessel (column width 153.5mm)

## Slide 29: Conclusion
Multi-Grid developed and characterized.
Large-Area detectors possible again
Operated at CNCS over 11 months.
Multi-Grid baseline detector for
    CSPEC and T-REX at ESS
Test at SEQUOIA proposed
Instrument-focused design is underway
Thanks to the colleagues at SNS for this opportunity and cooperation!
WP 4.3: Large-Area Detectors
Horizon 2020 grant agreement 676548
2017 JINST 12 P04030
(https://arxiv.org/abs/1703.03626 )

## Slide 30: Acknowledgements and Publications
ILL:
Bruno Guerard, Jean-Claude Buffet,
Jean-Francois Clergeau, Anthony Leandri ,
Victor Buridon, Fabien Lafont
ESS:
Anton Khaplanov, Richard Hall-Wilton, Oliver Kirstein, Tomasz Brys, Michail Anastasopoulos, Isaak Lopez Higuera, Carina Höglund*, Linda Robinson*
Centre for Energy Research (Hungary):
Eszter Dian
Linköping University:
Jens Birch, Lars Hultman, (also *)
SNS:
Ken Herwig, Georg Ehlers, Michelle Everett, Kevin Berry
Earlier – the participants of the CRISP project on Large-Area detectors.
Earlier publications:
B4C layers:
*C. Höglund et al, J of Appl. Phys. 111, 104908 (2012)
Characterization:
*A. Khaplanov et al., arXiv:1209.0566 (2012)
*B Guerard et al., NIMA, 720, 116-121 (2013), http://dx.doi.org/10.1016/j.nima.2012.12.021iJ
*J. Correa et al., Trans. Nucl. Sc. (2013), DOI: 10.1109/TNS.2012.2227798
*A. Khaplanov et al., (2014) J. Phys.: Conf. Ser. 528 012040 doi:10.1088/1742-6596/528/1/012040
Gamma sensitivity:
*A. Khaplanov et al., JINST 8, P10025 (2013), arXiv:1306.6247
Alpha background:
*A. Khaplanov et al., JINST 10, P10019 (2015); doi:10.1088/1748-0221/10/10/P10019
Latest publication:
A.Khaplanov et al. “Multi-Grid Detector for Neutron Spectroscopy: Results Obtained on Time-of-Flight Spectrometer CNCS" https://arxiv.org/abs/1703.03626 2017 JINST 12 P04030
WP 4.3: Large-Area Detectors
Horizon 2020 grant agreement 676548

## Slide 31: Extra slides

## Slide 32: IN5 Demonstrator – Large-Scale
Detector:
2.4 m2 active area
Blade x18432:
enriched B4C coating
good adhesion, uniformity,
Module x8:
wire-frame coincident readout
Grid x1024:
low activity,
minimal dead material
80cm
3 m

## Slide 33: Aluminium Bragg Edge
Incident n
Sample
Sample environment
Detector
Scattering effects from SE and detector

## Slide 34: Energy Reconstruction
Energy transfer: Ef-Ei
Raw ToF spectrum:
no information on geometry used
dE spectrum: Including information on position of each detected neutron.

## Slide 35: Single Crystal Reflection
He3
Multi-Grid
Image recorded in each detector as a function of time during elastic peak
Each frame advances 30 μs
No loss of position resolution or saturation observed in Multi-Grid

## Slide 36: C-SPEC
Neighbour vessels configuration, minimizing dead spaces