## Slide 1: Straw tubes for LoKI
Davide Raspino
IKON14
12-02-2018

## Slide 2: Outline
Multiplexing readout
  Method
  Performance
  Position corrections
Electronics
  New preamps
Mechanics
  Single unit
  On the instrument

## Slide 3: Multiplexing
n
1 m2 detector with straw tubes
5 layers
1400x2 preamp + ADC Chs
1 m2 with 8 mm ⊘ 3He
120x2 preamp + ADC Chs

## Slide 4: Multiplexing

## Slide 5: Position Resolution
BCS | Pos. Res. FWHM (mm) | Rate/BCS (kHz)
1 | 7.1 | 10.0
2 | 7.3 | 10.9
3 | 7.8 | 11.9
4 | 6.7 | 12.2
5 | 7.2 | 11.4
6 | 7.6 | 12.2
7 | 7.3 | 11.7
Cd
Cd

## Slide 6: Position Resolution
Cd
Cd
Cd

## Slide 7: Position Resolution
Cd
Cd
Cd

## Slide 8: Position Resolution vs Neutron rate
HT = 1000 V
HT = 1000 V

## Slide 9: Count Rate
Cd
Cd

## Slide 10: Count Rate Loss
~5% loss
every 200 kHz/1” tube
Detected Rate

## Slide 11: Position Corrections
Cd
Cd
Cd
Cd
Cd
Cd

## Slide 12: Position Corrections
Cd
Cd
Cd
Cd
Cd
Cd

## Slide 13: Position Corrections

## Slide 14: Position Corrections
Corrections vs Measured position
Corrected image
(mm)
(mm)

## Slide 15: Performance
Multiplexing 7x straw tubes reduce the number of Preamps/cables and ADCs by a factor 3.5
Position resolution of 8 mm FWHM up 300 kHz/1” tube and of 10 mm FWHM up to 500 MHz/1” tube (at 1000 V)
Count rate loss ~10% at 400 kHz/1” tube (detected rate)
Position correction very similar to the one needed for 3He tubes, few percent of the length of the tubes
All the shown position calculation used the position correction described

## Slide 16: Front End Electronics
16 chs preamp board
PCB size 115x38 mm2
Fully differential
Output connector  Ethernet
Noise lower than the ones use for the test shown here
Lower gain to allow to work at higher HT
Higher input impedance Position correction
Power dissipation: 220 mW – 14 mW/ch
Erik Schooneveld

## Slide 17: Mechanics
Anton Orszulik

## Slide 18: Straw tubes on LoKI
Detector Area (m2) | 9.09
Solid angle (sr) | 1.7
n. Panels | 9
Pack types | 3
Tube lengths (m) | 0.53,1,1.4
Federico Masi