## Slide 1: Beam Monitors
Fatima Issa, Steven Alcock , Richard Hall-Wilton, Anton Khaplanov, Scott Kolya, Vendula Maulerova ,
Thanks for the detector group at ISIS and SNS
Thanks to the instruments team for their inputs
Thanks to the manufacturers for their openness
www.europeanspallationsource.se
18-02-11

## Slide 2: Outline
Why Beam monitors?
Beam monitors market survey
Mechanical integration (BM on chopper)
Parasitic methods for monitoring the beam
Beam monitors at V20 at HZB
Beam monitor requirements per zone
Guidelines for Beam monitors at ESS
Summary
Notes: This talk will be an update to what has been presented during last IKON focusing on waht are main challenges per zone with their relative (potential) solutions
this slide follows on from what was presented at the last IKON. It is an update.

## Slide 3: ESS instruments
long-pulse and high neutron flux at ESS
Many advancements in the instrument performance
longer neutron guides, complex neutron optics and complex chopper systems

## Slide 4: Beam Monitor
Sample
source
neutron guide
chopper

## Slide 5: Beam Monitor
Sample
source
Sample
source
neutron guide
chopper
Beam monitor

## Slide 6: Beam Monitor
Transmitted beam
Incident beam
scattered
Sample
source
Beam monitor
neutron guide
chopper
Beam monitor
Beam monitor with low attenuation and high transmission factor
Window 1
Window 2

## Slide 7: Beam Monitor
Nuclear reactions used for thermal neutron detection
3He
14N
10B
Thickness 1cm at 1.8A
Beam monitors are thermal neutron detectors with relatively low efficiency which can vary from
10-6 to 10-1 depending on the instrument requirements
10B is in um in transmission mode, 14N and 3He is bar-cm thickness
Notes: Lookup table to know what is needed for a certain efficiency

## Slide 8: Types of Beam Monitor available today
MWPC from Mirrotron MWPC from Mirrotron Fission chamber from LND
MWPC from ORDELA GEM from CDT Scintillator from Quantum detector
https://indico.esss.lu.se/event/870/ https://indico.esss.lu.se/event/781/
Notes: What we have done so far we tested several types of BM from different vendor and we evaluted them. MWPC from …

## Slide 9: Beam Monitors-Gamma sensitivity measurement
Beam monitor
Gamma source
neutron source
Beam monitor
MWPC from Mirrotron
Fission chamber from LND
Gamma sensitivity studied with three gamma sources with low, medium, and high energy
source | Gamma energy KeV | Gamma intensity %
57Co | 122 | 85.6
 | 136 | 10.7
133Ba | 30-35 | 117
 | 80 | 36
 | 276 | 7
 | 303 | 18
 | 356 | 62
 | 384 | 9
60Co | 1173 | 100
 | 1332 | 100

## Slide 10: Beam Monitors- Neutron measurement at R2D2
Slit 1
Slit 2
Peak wavelength (Å)
400 2.00
911 0.878
711 1.12
511 1.54
822 0.943
311 2.41
933 0.804

## Slide 11: Beam monitors-main results
 | MWPC | MWPC | MWPC BM-100X50 | 2D-MWPC | 2D-GEM | Scintillator | Fission chamber
supplier | ORDELA | ORDELA | Mirrotron | Mirrotron | CDT | Quantum detector | LND
Isotope used for neutron capture | 3He | 14N | 3He | 3He | 10B | 6Li | 235U
Gas pressure mbar | Partial pressure 6,0795 | Partial pressure 81,06 | Partial pressure 6,5 | Partial pressure 0.4 | Total pressure 100 | ------ | Total pressure 1013,2
Filled gas | 3He+4He +CF4 | N+CF4 | 3He+CF4 | 3He+CF4 | Ar/CO2 | -------- | P10
Active Area (mm2) | 114 x 51 | 114 x 51 | 100 x 50 | 100 x 50 | Diameter 100 mm | 28 x 42 | Diameter 108.0 mm
Applied voltage (V) | 850 | 850 | 1300 | Anode at -3500V Drift at 1500V | -1000 | 650 | 300
Measured attenuation % at 2.4Å | 4.5 | 4.4 | 2.5 | 7.3 | 10.6 | 0.4 | 3.8
Calculated attenuation % at 2.4Å | 4.05 | 4.05 | 2.05 | 4.5 | 6.9 | 0.1 | 2.4
Measured Efficiency at 2.4Å | 1.2x10-3 | 2x10-5 | 1.1x10-3 | 1.5x10-4 | 3x10-2 | 5x10-3 | 10-4
Supplier Efficiency at 1.8 Å | 10-3 | 10-5 | 10-3 | 1.5x10-4 | 3x10-2 | 10-4 | 10-4
Scattering (%) | 3.9 ± 0.4 | 3.8±0.4 | 4±0.9 | 9±1 | 10.3±0.7 | 0.74±0.2 | 3.8±0.3
Beam monitor characterization  Customization to match specific requirements

## Slide 12
Horizon 2020 grant agreement 676548
https://journals.aps.org/prab/abstract/10.1103/PhysRevAccelBeams.20.092801

## Slide 13: Outline
Why Beam monitors?
Beam monitors market survey
Mechanical integration (BM on chopper)
Parasitic methods for monitoring the beam
Beam monitors at V20 at HZB
Beam monitor requirements per zone
Guidelines for Beam monitors at ESS
Summary

## Slide 14: Mechanical integration: Reminder
The chopper can run up to 22 Hz
BM at the backside of the chopper @Embla
** Thanks for Niko and Andrés from Chopper Group

## Slide 15: Mechanical integration: Reminder
The BM should be electrically isolated from the chopper.
Grounding is important follow grounding guidelines
ESS-0147271
Electrical contact  Significant noise
BM
Preamp
&shaper
MCA
** Thanks for Niko and Andrés from Chopper Group

## Slide 16: Looking for solutions …..
Gamma detector
Incident beam
In the Bunker area
In the Instrument Guideline
Ionization chamber form CDT
Fiber monitor used at ISIS
Vanadium foil
Gamma detector close to a chopper

## Slide 17: Looking for solutions …..
Electronics out of the beam
Low Mass Beam Monitor
Silicon detector with LiF in beam prototype on MIRACLES

## Slide 18: Outline
Why Beam monitors?
Beam monitors market survey
Mechanical integration (BM on chopper)
Parasitic methods for monitoring the beam
Beam monitors at V20 at HZB
Beam monitor requirements per zone
Guidelines for Beam monitors at ESS
Summary

## Slide 19: Beam monitor for chopper diagnostics
Fibre beam monitor used at LET at ISIS
Aluminum tube is about 3 mm in diameter.
Small rods of GS glass scintillator, each about 10 mm in length and 0.5 mm in diameter
**Info from Niko and Nigel and Robert Bewley and Daniel Pooley Thanks!
The fibres are about 0.5mm diameter and are protected by sitting in a thin Aluminium tube just a couple of mm across
to check chopper timing and that they are opening and closing exactly when they should be
Vacuum compatiable
The LET ‘fibre’ monitors are not off-the-shelf-options :
They are very fragile and difficult to make.
2) The cost of precision guide-apertures often outweighs the benefits realised.
3) parts/spares are not off-the-shelf

## Slide 20: Parasitic methods for monitoring the beam
Time resolution is not great, as currently implemented
as a white beam monitor, or to get phasing of a slow band width chopper is fine.
Detectors are placed away from the radiation environment
A project has just started with a master student, timescale 18 months ,
 Looking for improved geometry and simple engineering solutions:
make the foil and detector dimensions smaller  reduce time uncertainty with simplified engineering
Very preliminary Geant4 simulation
**Info from Nigel and Robert Bewley and daniel pooley Thanks!

## Slide 21: Parasitic methods for monitoring the beam
Using Gamma detector
Using the boron layer on the chopper disc  information about the chopper rotation  this is done using a gamma detector outside the beam.
Gamma detector
Incident neutron beam
energy= 0.48MeV
Notes: Another way is no matterial at all in the beam u take the advantage that ur chopper is already coated with B

## Slide 22: Parasitic methods for monitoring the beam
Boron carbide
LaBr scintillator detector from Saint-Gobain is surrounded by a mirrobor 2mm (80% B4C)
A gamma peak at 480 keV can be recognized in the presence of boron
The gamma peak at 480 keV can be recognized in the presence of different gamma sources
Notes: Very rough measurement was performed at SF using n sources to make sure this makes sense ..

## Slide 23: Parasitic methods for monitoring the beam at V20 HZB
A mini chopper coated with Boron used in the beam
LaBr gamma detector used for detecting the gamma peak from the 10B interaction
One beam monitor was placed before and another one after the mini chopper for the V20 integration test
** Thanks for Niko and Andrés from Chopper Group
ESS Mini chopper
LaBr gamma detector, no shielding, no optimization of scintillator or chopper

## Slide 24: Parasitic methods for monitoring the beam at V20 HZB
At V20 at HZB
A gamma peak at 480 keV can be recognized when beam is on
The peak intensity increases by ~ 3 times when chopper is closed  proof of concept
ESS will continue on this:
* Proof of concept paper by summer and Realistic performance/cost/engineering (timescale 12-18 months)
Very preliminary

## Slide 25: Outline
Why Beam monitors?
Beam monitors market survey
Mechanical integration (BM on chopper)
Parasitic methods for monitoring the beam
Beam monitors at V20 at HZB
Beam monitor requirements per zone
Guidelines for Beam monitors at ESS
Summary

## Slide 26: Beam monitor electronics
Front end
AMP
Event formation
Timing EPICS
DMSC
DAQ
Beam monitor
Work done by: Steven, Scott (DG) Tobias and Janos (DMSC)
example
Notes: So now what is a beam monitor as a system is the detector itself with electronics ..

## Slide 27: V20 test at HZB
Front end
AMP
Event formation
Timing EPICS
DMSC
DAQ
Beam monitor
DMSC and ICS
BM electronics several options (for more details see Scott talk at IKON13):
Integration Demonstrator (Done)
Talin
Talin ++
Standard detector readout

## Slide 28: V20 test at HZB
V20 test at HZB-Berlin:
Detector Group
Chopper Group
DMSC
ICS
For more details See Chopper group talk on
Wednesday 14 Feb at the Plenary session at 10:00
System integration shown
Successful test:
DMSC produced time-of-flight spectra based on the D.G. readout from the beam monitors,
and Diagnose chopper using two BMs

## Slide 29: Outline
Why Beam monitors?
Beam monitors market survey
Mechanical integration (BM on chopper)
Parasitic methods for monitoring the beam
Beam monitors at V20 at HZB
Beam monitor requirements per zone
Guidelines for Beam monitors at ESS
Summary

## Slide 30: Beam monitor system Integration
rarely
rarely
Mechanical
alignment

## Slide 31: Facility maintenance and operation Requirements
No access to the bunker or guides just for BM maintenance
Environmental requirement (rate, humidity, temperature,..)
T0 information
Diagnose flux lost
Flux along the beam line
BM system
Stability
Cross-Calibration
Radiation hardness
(lifetime >10years @ normal flux)
Reliability
Dynamic Range
Commissioning
Data mining
Notes: ..

## Slide 32: Beam monitor requirements
Instruments | Number of BM | Time resolution (μs) | Position resolution | Area (cm x cm) | Beam size mm x mm | Flux (n/cm2.s) | Pulse duration (ms) | Vacuum/air
LKOI | 2 or 3 for normalization | 100 | None | TBD | NM | NM | 53 to 63 | vacuum
 | 2 for diagnostics | 100 | None | TBD | NM | NM |  | vacuum
FREIA | 2 for normalization | Same as wavelength resolution | None | 2.5 x 4 | NM | NM |  | air
 | One behind each chopper pair for diagnostics | Same as chopper resolution | None | As small as non absorbing | NM | NM |  | vacuum
MIRACLES | 1 for normalization | 10 μs | None | 5x5 | Size of the final neutron guide | Highest flux on sample 5x106 |  | vacuum
 | 2 for diagnostics | 10 μs | None | up to 12x12 | Beam guide size |  |  | vacuum
VESPA | 1 for normalization | 15-20 μs | None | TBD | 35 x 35 | ….. | TBD | Vacuum (preferable)
 | 3 or4 for diagnostics | 50-100 μs | None | TBD | 35 x 35 to 50 x50 | …… | TBD | Vacuum (preferable)
HEIMDAL | 1or 2 for normalization | 10 μs | 0.2 mm x 0.2 mm | 4 x 4 | NM | 107 to 2x109 n/s/full area |  | air
 | 2 for diagnostics | 100 μs | None | Up to 10 x 10 cm | NM | 3x108 to 5x109 n/s/full area |  | vacuum
CSPEC | 1 for normalization | <2μs | 2 mm | 2.2 x 4 cm | 22 x 4 | 6x10 7 to 6x 109 n/s/Å/cm2 | NM | vacuum
 | 4 for diagnostics | 10 and 200 | 3mm | Up to 11.1 x 12 cm | 70 x 55 to 111x 120 | 6x10 7 to 6x 109 n/s/Å/cm2 | NM | vacuum
T-REX | 1 for normalization | 1 μs | 1x 1 mm | 1.5 x 3 | 15 x 30 | 1010 | 0.1 ms | vacuum
 | 2 for diagnostics | 2 μs | 1x 1 mm | 6 x 8.5 | 60 x 85 | 1010to 1011 | 1 ms to 0.5 ms | air
BEER | 1 for normalization | 10 μs | None | 50 x 100 mm² | Max 40x 80 | 107 to 109 |  | air
 | 1 for diagnostics | 10 μs | None | 25 x 85 mm² | 20 x 80 |  |  | Vacuum
BIFROST | 2 for normalization | 100 us | Optional | 5 x 5 | 40 x 40 | 107 to 1010 | Btw 0.1 and 3 ms | air
 | 2 for diagnostics | none | none | 4 x6 and 12x12 | 30x50 and 100x100 | 107 to 1012 | Btw 0.1 and 3 ms | air
NMX | None for normalization | non | Non |  | None | ……. |  | 
 | 3 for diagnostics | <1ms | 1-2 mm | 5 x 5 | 45 x 45 at choppers 30 x30 at end of guide | ………. |  | air
ESTIA | 0-1 for normalization | 1 μs | Not evaluated | > 6 x 6 | > 60x60 | <1010 n/s | Sufficient statistics needed 104 cts/bin | vacuum
 | 4 for diagnostics | 1 ms | < 1mm | 4 x4 and 2x4 | 20 x 15 and 20 x 30 | 1010 to 1012 n/s | No specific statistics needed | vacuum
ODIN | 1 for normalization | 100 | Maybe 1 mm | Max 10 x10 | 4 x 4 | 1011 | 60 to 120 ms | 
 | 2 for diagnostics | 100 | Maybe 1 mm | Max 10x 10 | 3 x 8 and 10x10 | 6x1010 to 3 x1010 | 20 to 40 and 35 to 70 | 
Notes: To start understanding we started by filling in this table very basic requirments like beam zize, flux, time resolution position resolution , vacuum or not …all based on this document u have seen this

## Slide 33: Beam monitor requirements
Time resolution : 1 μs to 1ms
Position resolution: 0.2mmx 0.2mm up to 3 mm
Flux on each monitor: 107 n/s/cm2 up to 1012 n/s/cm2
Beam size: 2x 15 mm up to 111x120 mm
Active area: 15x 30 mm up to 120x120 mm
Instruments with no Position sensitive BM: LOKI-FREIA-VESPA-BEER-BIFROST
Majority of BM are in vacuum: LOKI-MIRACLES- VISPA- CSPEC-ESTIA-ODIN
Main concern: attenuation, scattering, enough statistics per measurement
Preliminary
This is based on what we received from the instrument team
See Rasmus’ talk on BIFROST from last IKON  this is the level of details to be considered we need this level of details

## Slide 34: Beam monitors per zone
ODIN
DREAM
NMX
MIRACLES
BEER
C-SPEC
T-REX
MAGIC
BIFROST
HEIMDAL
FREIA
LoKI
SKADI
VESPA
ESTIA
Bunker area
Instrument Guideline
Close to the sample
Notes: To better understand the requirements for 15 instrumentswe have to have to divide this per zone and deal with it per zone. Instead of looking for all the requiremnts at one time because instrument have some similar challenges per zone like in the bunker, along the guide and close to the sample, similar challenges but not nesscacerly the same

## Slide 35: Beam monitors per zone
Neutron Flux at various locations
1010 - 1012 n/cm2.s
107- 1010 n/cm2.s
Beam stop
Bunker area
Guideline
Sample area
“Flux” BM
“Flux” BM
“Position sensitive” BM
“Position sensitive” Transmission BM
n-camera
attenuation
scattering
radiation hardness
Saturation
mechanical integration
attenuation
Challenges
scattering
F(t) F(t) F(x,y,t) F(t) or F(x,y,t)
F(x,y)
Flux (x,y,t)
Very different challenges and requirements

## Slide 36: Beam monitors per zone
Instrument | Bunker | Guideline | Sample | Camera | Transmission
LOKI | 0 | 2 | 2 | 1 | 2
C-SPEC | 1-2 | 2 | 1 | 1 | 1
T-REX | 0? | 2? | 1 | 0 | 0
NMX | 0 | 2 | 1 | 1 | 0
BEER | 1 | 0 | 1 | 1 | 0
HEIMDAL | 1 | 1 | 1-2 | 0 | 0
ODIN | 2 | 1 | 1 | 0 | 0
ESTIA | 2? | 1? | 0? | 1 | 0
FERIA | ? | ? | 1-2 | 0 | 0
BIFROST | 1 | 1 | 1-2 | 1 | 1
MIRACLES | 0 | 1 | 1 | 1 | 1
VESPA | 2 | 1 | 1 | 0 | 1
SKADI | 0 | 1 | 1 |  | 0
Preliminary based on
BM requirement document
and TG2 documents
Partially updated

## Slide 37: Beam monitor in the Bunker
Instrument | Bunker | Time resolution (us) | Position resolution (mm x mm) | Beam size (cm x cm) | Local Instantaneous Neutron flux (n/cm2.s)
LOKI | 0 |  |  |  | 
C-SPEC | 1-2 |  |  |  | 
T-REX | 0? |  |  |  | 
NMX | 0 | No | No | 3 x 4.5 | 1x1011
BEER | 1 |  |  |  | 
HEIMDAL | 1 |  |  |  | 
ODIN | 2 | 100 | No | 3 x 8 | 1.4 x1012
ESTIA | 2? |  |  |  | 
FERIA | ? |  |  |  | 
BIFROST | 1 |  |  |  | 
MIRACLES | 0 |  |  |  | 
VESPA | 2 | 10 | Yes | 4. x 4. | ~13 x1011
SKADI | 0 |  |  |  | 
Waiting for update

## Slide 38: Beam monitor in the Bunker
Challenges for this region:
radiation hardness, durability, maintainability, background (fast and gamma), Neutron flux
Purpose:
facility monitoring, flux monitoring in time, commissioning, data mining
Preliminary recommendation:
ionization chamber , GEM or fission chamber
Things to be solved/proven:
Radiation hardness, Flux and dynamic range
Disadvantages:
Accessibility,
Bunker area:
is the area with highest radiation:
fast neutron flux are the same intensity 1011 -109 n/s
Once you lose line of sight, the fast neutron fluxes are down to 10^3 or 10^4 / s.
80 MGray per year for Gamma
 1018 n/year.

## Slide 39: Beam monitor in the guideline
Instrument | Guideline | Time resolution (us) | Position resolution (mm x mm) | Beam size (cm x cm) | Local Instantaneous Neutron flux (n/cm2.s)
LOKI | 2 |  |  |  | 
C-SPEC | 2 |  |  |  | 
T-REX | 2? |  |  |  | 
NMX | 2 | <1ms | 1-2 mm | 4.5 x 4.5 | 5x1011 -1x109
BEER | 0 |  |  |  | 
HEIMDAL | 1 |  |  |  | 
ODIN | 1 | 100 | No | 10 x 10 | 7 x1011
ESTIA | 1? |  |  |  | 
FERIA | ? |  |  |  | 
BIFROST | 1 |  |  |  | 
MIRACLES | 1 |  |  |  | 
VESPA | 1 | 10 | Yes | 5.5 x 5.5 | ~8 x1010
SKADI | 1 | 1 | 1 x 1 | 3x3 | 3 x1010
Waiting for update

## Slide 40: Beam monitor in the guideline
challenges for this region:
attenuation, number of gaps, mechanical integration, cost
purpose:
facility monitoring, flux monitoring , chopper diagnostics and commissioning
our preliminary recommendation:
Passive monitors (Vd, gamma detector close to a chopper, fibre monitors as the one at LET,
Si detectors coated and LiF foil in the beam)
Things to be solved/proven:
Performance, mechanical integration(gaps), cost
Disadvantages:
time resolution, take beyond proof of concept to engineered item
The instrument guideline contains
Choppers, guideline section  constrains on mechanical integration

## Slide 41: Beam monitor near the sample
Instrument | Sample | Time resolution (us) | Position resolution (mm x mm) | Beam size (cm x cm) | Local Instantaneous Neutron flux (n/cm2.s)
LOKI | 2 |  |  |  | 
C-SPEC | 1 |  |  |  | 
T-REX | 1 |  |  |  | 
NMX | 1 | <1ms | 1-2mm | 3 x3 | 2X109
BEER | 1 |  |  |  | 
HEIMDAL | 1-2 |  |  |  | 
ODIN | 1 | 100 | No? | 4 x 4 | 7 x1012
ESTIA | 0? |  |  |  | 
FERIA | 1-2 |  |  |  | 
BIFROST | 1-2 |  |  |  | 
MIRACLES | 1 |  |  |  | 
VESPA | 1 | 10 | 3 x 3 | 3. X 3. | 3.5 x1010
SKADI | 1 | 1 | 1 x 1 | 1x 1 | 7 x109
Waiting for update

## Slide 42: Beam monitor near the sample
challenges for this region:
attenuation, mechanical integration, cost , PSD, scattering, specific material, wide range of wavelengths and fluxes
purpose:
facility monitoring, science, problem solving, alignment, (varied), normalization
our preliminary recommendation:
depends a lot on requirements, can be position sensitive or not
Very dependent upon individual needs of instrument+ use case per monitor
Things to be solved/proven:
Need a detailed set of requirement+ use case per monitor
Disadvantages …

## Slide 43: Beam monitor after the sample “Transmission monitor”
Instrument | Transmission | Time resolution (us) | Position resolution (mm x mm) | Beam size (cm x cm) | Local Instantaneous Neutron flux (n/cm2.s)
LOKI | 2 |  |  |  | 
C-SPEC | 1 |  |  |  | 
T-REX | 0 |  |  |  | 
NMX | 0 under consideration | <1ms | No | 0.5x 0.5 | 2X109
BEER | 0 |  |  |  | 
HEIMDAL | 0 |  |  |  | 
ODIN | No | No | No | ? | ?
ESTIA | 0 |  |  |  | 
FERIA | 0 |  |  |  | 
BIFROST | 1 |  |  |  | 
MIRACLES | 1 |  |  |  | 
VESPA | 1 | 10 | 3 x3 | 3. X 3. | 3.5 x1010
SKADI | 0 |  |  |  | 
Waiting for update

## Slide 44: Beam monitor after the sample “Transmission monitor”
challenges for this region:
Backscattering, radiation hardness
purpose:
Science , alignment, Transmitted flux on the sample
our preliminary recommendation:
depends a lot on requirements, might be the rear detector can handle this
Things to be solved/proven:
Away from sample
Disadvantages:
Backscattering to be eliminated

## Slide 45: Outline
Why Beam monitors?
Beam monitors market survey
Mechanical integration (BM on chopper)
Parasitic methods for monitoring the beam
Beam monitors at V20 at HZB
Beam monitor requirements per zone
Guidelines for Beam monitors at ESS
Summary

## Slide 46: Beam monitor Guidelines
Accessibility/reliability | Radiation hardness materials
Mounting | If along the chopper or the guide electrically isolated
Movement | Not preferable option. Not in bunker, prefer not in guide. Needs very good justification if needed Mounting and positional reliability should be considered and cable rotation has to be reviewed Continuity of diagnostics? Enable data mining
Gas/cooling | Probably BM as sealed system. Gas/cooling complicated to integrate Needs very good justification if used
Failure/maintenance | Whenever possible parasitic maintenance  component should be reviewed and qualified for reliability and availability Desire no access for BM maintenance only
Electronics | BM will have standard backend electronics for maintenance and integration reasons
Racks | Racks maybe separated from detector rack maybe per grounding zone
Data | Event mode expected as default mode
Fission chambers | Preliminary recommendation only in or around the bunker region not close to the sample because of fast neutron emission. Before losing line of sight to moderator and out of line of sight to sample. (main disadvantage of the fission chamber is that in addition to the desired fission products three fast MeV neutrons on average and MeV gamma radiation are produced during thermal fission) (flux of 1010 n/s  emission of 105 n/s fast neutrons)( assuming fission chamber efficiency 10-5) Use of 235U is allowed by sustainable of material.

## Slide 47: Minimum information to be collected
Activation /materials | Specific material as specified by who ? Magnetic or non magnetic ?
Location | Bunker /Guideline/near sample
Mounted | On chopper / on guide,..
Flux and beam size | Technology used
Mechanical constrains | Minimum afforded gap
Vibration | On a high speed chopper ??
Number of counts per bin | Efficiency/ accuracy of measurement?
Medium | Vacuum / air  thickness of window?
….. | …….

## Slide 48: Minimum information taken before monitor build…
Info from
ISIS Neutron Detector Group
Level of detail here much greater than the level of detail presently provided to us

## Slide 49: Beam monitors requirements versus Power ramp-up ?
Severe challenges: wide range of collimations and power ramp-up  Beam monitors with different requirements
Purchase many beam monitors ? Or adjusting beam monitor efficiency once per year not more??
Monitors needed at commissioning might be different to monitors needed at full power due to dynamic range?
Maybe we have some time to get this right.
Example from SNS
** info from Technical Specifications Beamline 4 Reflectometer Beam Monitors
2006: Beam monitor requirements range over six orders of magnitude in neutron count rate
2009: Beam monitor requirements range over nine orders of magnitude in neutron count rate

## Slide 50: Guidelines for Neutron Beam monitors at ESS
Expected by summer 2018
“Draft started”

## Slide 51: Summary
Market Survey on what is available
Mechanical integration No electrical contact between BM and choppers
A new project has been started (timescale 18 months) for simulating, testing and engineering the Vanadium monitor
Preliminary results from V20 test has been presented  Engineered solution on timescale of 18 months
BM in the bunker  Radiation hardness  GEM/ionization chamber/ fission chamber could be used
BM in the Guidelines  Low attenuation/ limited gap  Parasitic methods could be used
BM near sample  low attenuation, scattering, wide flux range  Depends a lot on each instrument requirements
High level of detailed information will be needed
For facility maintenance  Preliminary recommendations min 3 BM per instrument (Bunker, Guideline, sample)
Guidelines for neutron beam monitors at ESS will be available in summer 2018

## Slide 52
THANK YOU
Special thanks to :
Detector group at ISIS ( Nigel and Robert Bewley and Daniel Pooley) and SNS (K. Berry)
Chopper Group at ESS ( Niko and Andrés)
Instruments team for their inputs