## Slide 1: Ion Source and fence designed for it
Øystein Midttun
Ion source and LEBT system leader
www.europeanspallationsource.se
8 February 2018

## Slide 2: The ESS proton source is a microwave discharge ion source
2
Waveguide
Solenoid magnets
Plasma chamber
H2 gas injection
High voltage insulator
Extraction aperture
Extraction electrodes
Pumping ports
50 cm

## Slide 3: Picture from the inside of the high voltage platform

## Slide 4: Ion source and low energy beam transport

## Slide 5: We have designed a high voltage safety fence with two main functions
1: protect against the 75 kV of the high voltage platform
75 kV

## Slide 6: We have designed a high voltage safety fence with two main functions
2: protect against x-ray coming from the ion source during beam extraction
fig 5
Report: PS-ESS radioprotection report 30.03.17, ESS-0114736
Normal operation:
< 15 µSv/h
Maximum observed:
380 µSv/h
Radiation level below
natural background
Measured dose rate levels
above natural background
Measurements at ion source test stand, INFN-LNS
2 mm lead shielding around the ion source

## Slide 7: X-rays were only measured with the beam on
The x-rays originate from back streaming electrons (up to 75 keV) from the beam line to the ion source
Plasma alone is not sufficient to produce a measurable amount of x-rays
No beam -> no x-rays
No high voltage -> no beam

## Slide 8: Our goal is to safely run the ion source while there is access to the accelerator tunnel
Ion source commissioning can take place at the same time as installation (of other stuff)
No access control to the tunnel needed
PSS0 will ensure that there no one can approach the HV platform while it’s on
By lead shielding the HV safety fence, the area outside can be classified as a public radiation zone

## Slide 9: 3D model of high voltage safety fence
Floor-to-ceiling walls
2.5 m high, 2 mm thick
lead shielding
1 m high stainless steel mesh (2 mm spacing, IP3x)
Cooling water inlet
for the ion source
Additional lead shielding
around the LEBT

## Slide 10: 3D model of high voltage safety fence
Safety door
PSS0
High voltage cable
Isolated 400 VAC
Fibre optic control cables

## Slide 11: Top view of high voltage safety fence
Entrance
Connection to grounding
mesh in the floor
Ion source
water cooling
Connection to
wall grounding

## Slide 12: Ion Source and fence designed for it
Description of the HV protection cage: ESS-0122281